TITLE 1. ADMINISTRATION

DEPARTMENT OF GENERAL SERVICES

Final Regulation

Titles of Regulations: 1VAC30-45. Certification for Noncommercial Environmental Laboratories (adding 1VAC30-45-10 through 1VAC30-45-860).

1VAC30-46. Accreditation for Commercial Environmental Laboratories (adding 1VAC30-46-10 through 1VAC30-46-210).

Statutory Authority: §2.2-1105 of the Code of Virginia.

Effective Date: October 1, 2008.

Agency Contact: Nancy S. Saylor, Consultant to the Division of Consolidated Laboratory Services, Department of General Services, 600 North 5th Street, Richmond, VA 23219, telephone (804) 231-7980, FAX (804) 231-7980 or email nssaylor@erols.com.

Summary:

The regulations establish the certification program required by §2.2-1105 of the Code of Virginia for environmental laboratories submitting data to the Department of Environmental Quality under the state’s air, water and waste laws. There are two regulations, one for noncommercial environmental laboratories (1VAC30-45) and one for commercial environmental laboratories (1VAC30-46). Each regulation is organized into two parts. Part I of each regulation contains the provisions pertaining to the administration of the program. This part describes the process that owners or operators of environmental laboratories must use to be certified and to maintain certification under the program. Part II of each regulation contains the quality assurance and quality control standards that environmental laboratories must meet to be certified under the program. The standards in Part II of Chapter 45 have been developed for Virginia noncommercial environmental laboratories. The standards in Part II of Chapter 46 are the 2002 National Environmental Laboratory Accreditation Conference standards, which are incorporated by reference into the regulation.

Substantive changes since the proposed are as follows:

1. In Chapter 45, added specific quality control requirements for different types of testing (in lieu of generic requirements);

2. In Chapters 45 and 46, revised the fees;

3. In Chapter 45, substituted matrix, technology/method, and analyte/analyte group in the definition of field of testing for program, method and analyte;

4. In Chapter 46, limited certification to 12 months instead of 24 months;

5. In Chapter 46, substituted the 2003 NELAC standards for the 2002 NELAC standards; and

6. Added one year to the establishment of the program (when laboratories must be certified to submit data to DEQ).

Summary of Public Comments and Agency's Response: A summary of comments made by the public and the agency's response may be obtained from the promulgating agency or viewed at the office of the Registrar of Regulations.

CHAPTER 45 
CERTIFICATION FOR NONCOMMERCIAL ENVIRONMENTAL LABORATORIES

Part I
General Provisions

1VAC30-45-10. Purpose.

Section 2.2-1105 of the Code of Virginia directs the Division of Consolidated Laboratory Services to establish a program to certify environmental laboratories that perform tests, analyses, measurements or monitoring required [ by pursuant to ] the Commonwealth’s air, waste and water laws and regulations. This chapter sets out the required standards and the process by which owners [ or operators ] of noncommercial environmental laboratories may obtain certification for their laboratories. 1VAC30-46 covers commercial environmental laboratories and [ NELAC-accredited NELAP-accredited ] environmental laboratories [ located outside seeking reciprocal accreditation in ] Virginia.

1VAC30-45-20. Establishment of certification program.

A. Once the certification program has been established, laboratory certification shall be required before any environmental analyses performed by a noncommercial environmental laboratory may be used for the purposes of the Virginia Air Pollution Control Law, the Virginia Waste Management Act or the State Water Control Law (§10.1-1300 et seq., §10.1-1400 et seq., and §62.1-44.2 et seq., respectively of the Code of Virginia).

B. The certification program shall be established on [ the first day of the 25th month following the effective date of this chapter October 1, 2011 ].

1VAC30-45-30. Applicability.

A. This chapter applies to any owner [ or operator ] of a noncommercial environmental laboratory.

B. Any environmental laboratory owned by an agency of the federal government may be certified as follows:

1. [ By ] DGS-DCLS to the standards set out in this chapter, or

2. [ A By a ] federal primary accrediting authority to the standards established by the National Environmental Laboratory Accreditation Conference.

1VAC30-45-40. Definitions.

[ Where a term is defined in this section, the term shall have no other meaning, even if it is defined differently in the Code of Virginia or another regulation of the Virginia Administrative Code. Unless specifically defined in this section, the terms used in this chapter shall have the meanings commonly ascribed to them by recognized authorities. ]

"Acceptance criteria" means specified limits placed on characteristics of an item, process, or service defined in requirement documents.

"Accuracy" means the degree of agreement between an observed value and an accepted reference value.  Accuracy includes a combination of random error (precision) and systematic error (bias) components that are due to sampling and analytical operations.  Accuracy is an indicator of data quality.

[ "Algae" means simple single-celled, colonial, or multicelled, mostly aquatic plants, containing chlorophyll and lacking roots, stems and leaves that are either suspended in water (phytoplankton) or attached to rocks and other substrates (periphyton). ]

"Aliquot" means [ something that is contained an exact number of times in another, such as aliquot samples for testing or analysis a portion of a sample taken for analysis ].

"Analyte" means the substance or physical property to be determined in samples examined.

[ "Analytical batch" means a batch composed of prepared environmental samples (extracts, digestates or concentrates) that are analyzed together as a group.  An analytical batch can include prepared samples originating from various environmental matrices and can exceed 20 samples. ]

"Analytical method" means a technical procedure for providing analysis of a sample, defined by a body such as the Environmental Protection Agency or the American Society for Testing and Materials, that may not include the sample preparation method.

"Assessment" means the evaluation process used to measure or establish the performance, effectiveness, and conformance of an organization and its systems or both to defined criteria.

"Assessor" means the person who performs on-site assessments of laboratories’ capability and capacity for meeting the requirements under this chapter by examining the records and other physical evidence for each one of the tests for which certification has been requested.

"Audit" means a systematic evaluation to determine the conformance to quantitative and qualitative specifications of some operational function or activity.

"Authority" means, in the context of a governmental body or local government, an authority created under the provisions of the Virginia Water and Waste Authorities Act, Chapter 51 (§15.2-5100 et seq.) of Title 15.2 of the Code of Virginia.

"Batch" means environmental samples that are prepared together or analyzed together or both with the same process and personnel, using the same lot or lots of reagents.   [ See "analytical batch." "Analytical batch" means a batch composed of prepared environmental samples (extracts, digestates or concentrates) that are analyzed together as a group. An analytical batch can include prepared samples originating from various environmental matrices and can exceed 20 samples. "Preparation batch" means a batch composed of one to 20 environmental samples of the same matrix that meets the criteria in this definition for "batch" and with a maximum time between the start of processing of the first and last sample in the batch to be 24 hours.

"Benthic macroinvertebrates" means bottom dwelling animals without backbones that live at least part of their life cycles within or upon available substrates within a body of water. ]

"Blank" means a sample that has not been exposed to the analyzed sample stream in order to monitor contamination during sampling, transport, storage or analysis.  The blank is subjected to the usual analytical and measurement process to establish a zero baseline or background value [ and is sometimes used to adjust or correct routine analytical results. Blanks include the following types:

1. Field blank. A blank prepared in the field by filling a clean container with pure deionized water and appropriate preservative, if any, for the specific sampling activity being undertaken.

2. Method blank. A sample of a matrix similar to the batch of associated samples (when available) that is free from the analytes of interest and is processed simultaneously with and under the same conditions as samples through all steps of the analytical procedures, and in which no target analytes or interferences are present at concentrations that impact the analytical results for sample analyses ].

"Calibration" means to determine, by measurement or comparison with a standard, the correct value of each scale reading on a meter, instrument or other device.  The levels of the applied calibration standard should bracket the range of planned or expected sample measurements.

"Calibration curve" means the graphical relationship between the known values, such as concentrations, of a series of calibration standards and their instrument response.

"Calibration standard" means a substance or reference material used to calibrate an instrument.

[ "Certified reference material" means a reference material one or more of whose property values are certified by a technically valid procedure, accompanied by or traceable to a certificate or other documentation that is issued by a certifying body. ]

"Commercial environmental laboratory" means an environmental laboratory where environmental analysis is performed for another person.

"Corrective action" means the action taken to eliminate the causes of an existing nonconformity, defect or other undesirable situation in order to prevent recurrence.

"DGS-DCLS" means the Division of Consolidated Laboratory Services of the Department of General Services.

"Demonstration of capability" means the procedure to establish the ability of the analyst to generate data of acceptable accuracy and precision.

"Detection limit" means the lowest concentration or amount of the target analyte that can be determined to be different from zero by a single measurement at a stated degree of confidence.

"Environmental analysis" or "environmental analyses" means any test, analysis, measurement, or monitoring used for the purposes of the Virginia Air Pollution Control Law, the Virginia Waste Management Act or the State Water Control Law (§10.1-1300 et seq., §10.1-1400 et seq., and §62.1-44.2 et seq., respectively, of the Code of Virginia). For the purposes of these regulations, any test, analysis, measurement, or monitoring required [ by pursuant to ] the regulations promulgated under these three laws, or by any permit or order issued under the authority of any of these laws or regulations is "used for the purposes" of these laws.  The term shall not include the following:

1. Sampling of  [ air, water or waste water, solid and chemical materials, biological tissue, or air and emissions ].

2. Field testing and measurement of [ air, water or waste water, solid and chemical materials, biological tissue, or air and emissions ], except when performed in an environmental laboratory rather than at the site where the sample was taken.

[ 3. Taxonomic identification of samples for which there is no national accreditation standard such as algae, benthic macroinvertebrates, macrophytes, vertebrates and zooplankton. ]

"Environmental laboratory" or "laboratory" means a facility or a defined area within a facility where environmental analysis is performed. [ A structure built solely to shelter field personnel and equipment from inclement weather shall not be considered an environmental laboratory. ]

"Establishment date" means the date set for the accreditation program under 1VAC30-46 and the certification program to be established under this chapter.

"Establishment of certification program" or "established program" means that DGS-DCLS has completed the initial accreditation of environmental laboratories covered by 1VAC30-46 and the initial certification of environmental laboratories covered by 1VAC30-45.

"Facility" means something that is built or installed to serve a particular function.

"Field of [ testing" certification" ] means an approach to certifying laboratories by [ program matrix ], [ method technology/method ] and [ analyte analyte/analyte group ].

"Field testing and measurement" means any of the following:

1. Any test for parameters under 40 CFR Part 136 for which the holding time indicated for the sample requires immediate analysis; or

2. Any test defined as a field test in federal regulation.

The following is a limited list of currently recognized field tests or measures that is not intended to be inclusive:  continuous emissions monitoring; on-line monitoring; flow monitoring; tests for pH, residual chlorine, temperature and dissolved oxygen; and field analysis for soil gas.

"Finding" means an assessment conclusion that identifies a condition having a significant effect on an item or activity. An assessment finding is normally a deficiency and is normally accompanied by specific examples of the observed condition.

"Governmental body" means any department, agency, bureau, authority, or district of the United States government, of the government of the Commonwealth of Virginia, or of any local government within the Commonwealth of Virginia.

"Holding time (or maximum allowable holding time)" means the maximum time that a sample may be held prior to analysis and still be considered valid or not compromised.

"Initial certification period" means the period during which DGS-DCLS is accepting and processing applications for the first time under this chapter as specified in 1VAC30-45-60.

[ "International System of Units (SI)" means the coherent system of units adopted and recommended by the General Conference on Weights and Measures.

"Laboratory control sample" or "LCS" means a sample matrix, free from the analytes of interest, spiked with verified known amounts of analytes or a material containing known and verified amounts of analytes. It is generally used to establish intra-laboratory or analyst specific precision and bias or to assess the performance of all or a portion of the measurement system. "Laboratory control sample" or "LCS" may also be named laboratory fortified blank, spiked blank, or QC check sample. ]

"Laboratory manager" means the person who has overall responsibility for the technical operation of the environmental laboratory and who exercises actual day-to-day supervision of laboratory operation for the appropriate fields of testing and reporting of results.  The title of this person may include but is not limited to laboratory director, technical director, laboratory supervisor or laboratory manager.

"Legal entity" means an entity, other than a natural person, who has sufficient existence in legal contemplation that it can function legally, be sued or sue and make decisions through agents as in the case of corporations.

[ "Limit of detection" or "LOD" means an estimate of the minimum amount of a substance that an analytical process can reliably detect. An LOD is analyte and matrix specific and may be laboratory dependent.

"Limit of quantitation" or "LOQ" means the minimum levels, concentrations, or quantities of a target variable (e.g., target analyte) that can be reported with a specified degree of confidence. ]

"Local government" means a municipality (city or town), county, sanitation district, or authority.

[ "Macrophytes" means any aquatic or terrestrial plant species that can be identified and observed with the eye, unaided by magnification. ]

"Matrix" means the component or substrate that may contain the analyte of interest. [ For purposes of batch and quality control requirement determinations, the following matrix types shall be used: A matrix can be a field of certification matrix or a quality system matrix. ]

1. [ Drinking Field of certification matrix. These matrix definitions shall be used when certifying a laboratory.

a. Nonpotable ] water. Any aqueous sample that has [ not ] been designated a potable or potential potable water source.  [ 2. Nonpotable water. Any aqueous sample excluded from the definition of drinking water matrix. ] Includes surface water, groundwater, effluents, water treatment chemicals, and TCLP or other extracts.

[ 3. b. ] Solid and chemical materials. Includes soils, sediments, sludges, products and byproducts of an industrial process that results in a matrix not previously defined.

[ 4. c. ] Biological tissue. Any sample of a biological origin such as fish tissue, shellfish, or plant material. Such samples shall be grouped according to origin  [ , i.e., by species ] .

[ 5. d. ] Air and emissions. Whole gas or vapor samples including those contained in flexible or rigid wall containers and the extracted concentrated analytes of interest from a gas or vapor that are collected with a sorbent tube, impinger solution, filter or other device.

[ 2. Quality system matrix. For purposes of batch and quality control requirement determinations, the following matrix types shall be used:

a. Drinking water.  Any aqueous sample that has been designated a potable or potential potable water source.

b. Aqueous. Any aqueous sample excluded from the definition of drinking water matrix or saline/estuarine source. Includes surface water, groundwater, effluents, and TCLP or other extracts.

c. Saline/estuarine. Any aqueous sample from an ocean or estuary, or other salt water source.

[ d. Nonageous liquid. Any organic liquid with less than 15% settleable solids. ]

e. Biological tissue. Any sample of a biological origin such as fish tissue, shellfish, or plant material. Such samples shall be grouped according to origin.

f. Solids. Includes soils, sediments, sludges and other matrices with more than 15% settleable solids.

g. Chemical waste. A product or by-product of an industrial process that results in a matrix not previously defined.

h. Air and emissions.  Whole gas or vapor samples including those contained in flexible or rigid wall containers and the extracted concentrated analytes of interest from a gas or vapor that are collected with a sorbent tube, impinger solution, filter or other device.

[ "Matrix spike (spiked sample or fortified sample)" means a sample prepared by adding a known mass of target analyte to a specified amount of matrix sample for which an independent estimate of target analyte concentration is available. Matrix spikes are used, for example, to determine the effect of the matrix on a method’s recovery efficiency.

"Matrix spike duplicate (spiked sample or fortified sample duplicate)" means a second replicate matrix spike prepared in the laboratory and analyzed to obtain a measure of the precision of the recovery for each analyte. ]

"National Environmental Laboratory Accreditation Conference (NELAC)" means a voluntary organization of state and federal environmental officials and interest groups with the primary purpose to establish mutually acceptable standards for accrediting environmental laboratories.  A subset of NELAP.

"National Environmental Laboratory Accreditation Program (NELAP)" means the overall National Environmental Laboratory Accreditation Program of which NELAC is a part.

"National Institute of Standards and Technology" or "NIST" means an agency of the U.S. Department of Commerce’s Technology Administration that is working with EPA, states, NELAC, and other public and commercial entities to establish a system under which private sector companies and interested states can be certified by NIST to provide NIST-traceable proficiency testing (PT) samples [ to those laboratories testing drinking water and wastewater ].

"Negative control" means measures taken to ensure that a test, its components, or the environment do not cause undesired effects, or produce incorrect test results.

"Noncommercial environmental laboratory" means either of the following:

1. An environmental laboratory where environmental analysis is performed solely for the owner of the laboratory.

2. An environmental laboratory where the only performance of environmental analysis for another person is one of the following:

a. Environmental analysis performed by an environmental laboratory owned by a local government for an owner [ or operator ] of a small [ sewage wastewater ] treatment [ plant system ] treating domestic sewage at a flow rate of less than or equal to 1,000 gallons per day.

b. Environmental analysis performed by an environmental laboratory operated by a corporation as part of a general contract issued by a local government to operate and maintain a [ sewage wastewater ] treatment [ facility system ] or a waterworks.

c. Environmental analysis performed by an environmental laboratory owned by a corporation as part of the prequalification process [ for or to confirm the identity or characteristics of material supplied by ] a potential [ or existing ] customer [ or generator ] as required by a hazardous waste management permit under 9VAC20-60.

d. Environmental analysis performed by an environmental laboratory owned by a Publicly Owned Treatment Works (POTW) for an industrial source of wastewater under a permit issued by the POTW to the industrial source as part of the requirements of a pretreatment program under Part VII (9VAC25-31-730 et seq.) of 9VAC25-31.

e. Environmental analysis performed by an environmental laboratory owned by a county authority for any municipality within the county’s geographic jurisdiction when the environmental analysis pertains solely to the purpose for which the authority was created.

f. Environmental analysis performed by an environmental laboratory owned by an authority or a sanitation district for any participating local government of the authority or sanitation district when the environmental analysis pertains solely to the purpose for which the authority or sanitation district was created.

"Owner" [ or "operator" ] means any person who owns [ or , ] operates [ , leases or controls ] an environmental laboratory.

"Person" means an individual, corporation, partnership, association, company, business, trust, joint venture or other legal entity.

"Physical," for the purposes of fee test categories, means the tests to determine the physical properties of a sample.  Tests for solids, turbidity and color are examples of physical tests.

"Positive control" means measures taken to ensure that a test or its components are working properly and producing correct or expected results from positive test subjects.

"Precision" means the degree to which a set of observations or measurements of the same property, obtained under similar conditions, conform to themselves.  Precision is an indicator of data quality.  Precision is expressed usually as standard deviation, variance or range, in either absolute or relative terms.

"Primary accrediting authority" means the agency or department designated at the territory, state or federal level as the recognized authority with the responsibility and accountability for granting NELAC accreditation to a specific laboratory for a specific field of accreditation.

"Proficiency test or testing (PT)" means evaluating a laboratory’s performance under controlled conditions relative to a given set of criteria through analysis of unknown samples provided by an external source.

"Proficiency test (PT) field of testing" means the approach to offer proficiency testing by [ regulatory or environmental program, matrix type, and analyte maxtrix, technology/method, and analyte/analyte group ].

"Proficiency test (PT) sample" means a sample, the composition of which is unknown to [ both ] the analyst [ and the laboratory ], provided to test whether the analyst or laboratory or both can produce analytical results within specified acceptance criteria.

"Proficiency testing (PT) program" means the aggregate of providing rigorously controlled and standardized environmental samples to a laboratory for analysis, reporting of results, statistical evaluation of the results and the collective demographics and results summary of all participating laboratories.

"Program," in the context of [ field testing or a ] regulatory program, means the relevant U.S. Environmental Protection Agency program such as the water program under the Clean Water Act (CWA), the air program under the Clean Air Act (CAA), the waste program under the Comprehensive Environmental Response, Compensation and Liability Act (CERCLA or Superfund) or the waste program under the Resource Conservation and Recovery Act (RCRA).

"Publicly Owned Treatment Works (POTW)" means a treatment works as defined by §212 of the CWA, which is owned by a state or municipality (as defined by §502(4) of the CWA).  This definition includes any devices and systems used in the storage, treatment, recycling, and reclamation of municipal sewage or industrial wastes of a liquid nature. It also includes sewers, pipes, and other conveyances only if they convey wastewater to a POTW treatment plant. The term also means the municipality as defined in §502(4) of the CWA, which has jurisdiction over the indirect discharges to and the discharges from such a treatment works.

"Quality assurance" means an integrated system of activities involving planning, quality control, quality assessment, reporting and quality improvement to ensure that a product or service meets defined standards of quality with a stated level of confidence.

"Quality assurance officer" means the person who has responsibility for the quality system and its implementation.  Where staffing is limited, the quality assurance officer may also be the [ technical director laboratory manager ].

"Quality control" means the overall system of technical activities whose purpose is to measure and control the quality of a product or service so that it meets the needs of users.

"Quality manual" means a document stating the management policies, objectives, principles, organizational structure and authority, responsibilities, accountability, and implementation of an agency, organization, or laboratory, to ensure the quality of its product and the utility of its product to its users.

"Quality system" means a structured and documented management system describing the policies, objectives, principles, organizational authority, responsibilities, accountability, and implementation plan of an organization for ensuring quality in its work processes, products (items), and services.  The quality system provides the framework for planning, implementing, and assessing work performed by the organization and for carrying out required quality assurance and quality control.

"Range" means the difference between the minimum and maximum of a set of values.

"Reference material" means  a material or substance one or more properties of which are sufficiently well established to be used for the calibration of an apparatus, the assessment of a measurement test method, or for assigning values to materials.

[ "Reference standard" means a standard, generally of the highest metrological quality available at a given location, from which measurements made at that location are derived. ]

"Responsible official" means one of the following, as appropriate:

1. If the laboratory is owned or operated by a private corporation, "responsible official" means (i) a president, secretary, treasurer, or a vice-president of the corporation in charge of a principal business function, or any other person who performs similar policy-making or decision-making functions for the corporation or (ii) the manager of one or more manufacturing, production, or operating facilities employing more than 250 persons or having gross annual sales or expenditures exceeding $25 million (in second-quarter 1980 dollars), if authority to sign documents has been assigned or delegated in accordance with corporate procedures.

2. If the laboratory is owned or operated by a partnership, association, or a sole proprietor, "responsible official" means a general partner, officer of the association, or the proprietor, respectively.

3. If the laboratory is owned or operated by a governmental body, "responsible official" means a director or highest official appointed or designated to oversee the operation and performance of the activities of the environmental laboratory.

4. Any person designated as the responsible official by an individual described in subdivision 1, 2 or 3 of this definition, provided the designation is in writing, the designation specifies an individual or position with responsibility for the overall operation of the environmental laboratory, and the designation is submitted to DGS-DCLS.

"Sampling" means the act of collection for the purpose of analysis.

"Sanitation district" means a sanitation district created under the provisions of Chapters 3 (§21-141 et seq.) through 5 (§21-291 et seq.) of Title 21 of the Code of Virginia.

"Sewage" means the water-carried human wastes from residences, buildings, industrial establishments or other places together with such industrial wastes and underground, surface, storm, or other water as may be present.

"Simple test procedures" means any of the following:

1. Field testing and measurement performed in an environmental laboratory.

2. The test procedures to determine:

a. Biochemical oxygen demand (BOD),

b. Fecal coliform,

c. Total coliform,

d. Fecal streptococci,

e. [ E. coli,

f. Enterococci,

g. ] Settleable solids (SS),

[ f. h. ] Total dissolved solids (TDS),

[ g. i. ] Total solids (TS),

[ h. j. ] Total suspended solids (TSS),

[ i. k. ] Total volatile solids (TVS), and

[ j. l. ] Total volatile suspended solids (TVSS).

"Standard operating procedure (SOP)" means a written document that details the method of an operation, analysis or action whose techniques and procedures are thoroughly prescribed and which is accepted as the method for performing certain routine or repetitive tasks.

[ Standardized reference material (SRM)" means a certified reference material produced by the U.S. National Institute of Standards and Technology or other equivalent organization and characterized for absolute content, independent of analytical method.

"System laboratory" means a noncommercial laboratory that analyzes samples from multiple facilities having the same owner. ]

"TCLP" or "toxicity characteristic leachate procedure" means Test Method 1311 in "Test Methods for Evaluating Solid Waste, Physical/Chemical Methods," EPA Publication SW-846, as incorporated by reference in 40 CFR 260.11.  This method is used to determine whether a solid waste exhibits the characteristic of toxicity (see 40 CFR 261.24).

"Test" means a technical operation that consists of the determination of one or more characteristics or performance of a given product, material, equipment, organism, physical phenomenon, process or service according to a specified procedure.

"Test, analysis, measurement or monitoring required [ by pursuant to ] the Virginia Air Pollution Control Law" means any method of analysis required by the Virginia Air Pollution Control Law (§10.1-1300 et seq.); by the regulations promulgated under this law (9VAC5) [ , ] including any method of analysis listed either in the definition of "reference method" in 9VAC5-10-20, or listed or adopted by reference in [ 9VAC5-30, 9VAC5-40, 9VAC5-50 or 9VAC5-60 9VAC5 ] ; or by any permit or order issued under and in accordance with this law and these regulations.

"Test, analysis, measurement or monitoring required [ by pursuant to ] the Virginia Waste Management Act" means any method of analysis required by the Virginia Waste Management Act (§10.1-1400 et seq.); by the regulations promulgated under this law (9VAC20), including any method of analysis listed or adopted by reference in [ 9VAC20-60, 9VAC20-80, 9VAC20-101, or 9VAC20-120 9VAC20 ] ; or by any permit or order issued under and in accordance with this law and these regulations.

"Test, analysis, measurement or monitoring required [ by pursuant to ] the Virginia Water Control Law" means any method of analysis required by the Virginia Water Control Law (§62.1-44.2 et seq.); by the regulations promulgated under this law (9VAC25), including any method of analysis listed or adopted by reference in [ 9VAC25-31, 9VAC25-32, 9VAC25-110, 9VAC25-120, 9VAC25-151, 9VAC25-180, 9VAC25-190, 9VAC25-192, or 9VAC25-210 9VAC25 ]; or by any permit or order issued under and in accordance with this law and these regulations.

"Test method" means an adoption of a scientific technique for [ performing ] a specific measurement [ problem ], as documented in a laboratory standard operating procedure or [ as ] published by a recognized authority.

"Traceability" means the property of a result of a measurement whereby it can be related to appropriate standards, generally international or national standards, through an unbroken chain of comparisons.

"U.S. Environmental Protection Agency" means the federal government agency with responsibility for protecting, safeguarding and improving the natural environment (i.e., air, water and land) upon which human life depends.

"Virginia Air Pollution Control Law" means [ Chapter 13 ] §10.1-1300 [ et seq. ] of the Code of Virginia, which is titled "Air Pollution Control Board."

[ "Wastewater" means liquid and water-carried industrial wastes and domestic sewage from residential dwellings, commercial buildings, industrial and manufacturing facilities and institutions. ]

"Waterworks" means each system of structures and appliances used in connection with the collection, storage, purification, and treatment of water for drinking or domestic use and the distribution thereof to the public, except distribution piping.

[ "Zooplankton" means microscopic animals that float freely with voluntary movement in a body of water. ]

1VAC30-45-50. Scope of certification.

A. Noncommercial environmental laboratories shall be certified based on the general laboratory standards set out in Part II (1VAC30-45-200 et seq.) of this chapter and on the specific test methods or analysis, monitoring or measurement required by regulatory permit or other requirement under the Virginia Air Pollution Control Law, Virginia Waste Management Act or Virginia Water Control Law, the regulations promulgated under these laws, and by permits and orders issued under and in accordance with these laws or regulations.

B. DGS-DCLS shall review alternative test methods and procedures for certification when these are proposed by the applicant laboratory.  The provisions of 1VAC30-45-70 E and 1VAC30-45-90 B govern alternative test methods and procedures.

C. Certification shall be granted for [ a specific field or one or more ] fields of [ testing certification ], including the [ matrix, the ] technology and methods used by the noncommercial environmental laboratory, and the individual analytes or analyte groups determined by the particular method.

1VAC30-45-60. General: certification requirements.

A. Components of certification.  The components of certification include review of personnel qualifications, on-site assessment, proficiency testing, and quality systems.  The criteria for these components, set out in Part II (1VAC30-45-200 et seq.) of this chapter, shall be fulfilled for certification.

B. Individual laboratory sites and mobile laboratories.

1. Individual laboratory sites are subject to the same application process, assessments, and other requirements as environmental laboratories.  Any remote laboratory sites are considered separate sites and subject to separate on-site assessments.

2. Laboratories located at the same physical location shall be considered an individual laboratory site if these laboratories are owned [ or operated ] by the same person, and have the same laboratory manager and quality system.

3. [ Laboratories located at separate, noncontiguous physical locations may request to be considered as an individual laboratory site if these laboratories are owned by the same person and have the same laboratory manager and quality system. ]

[ 4. ] A mobile laboratory, which is configured with equipment to perform analyses, whether associated with a fixed-based laboratory or not, is considered an environmental laboratory and shall require separate certification.  This certification shall remain with the mobile laboratory and be site independent.  Moving the configured mobile laboratory to a different site will not require a new or separate certification.  Before performing analyses at each new site, the laboratory shall ensure that instruments and equipment have been checked for performance and have been calibrated.

1VAC30-45-70. Process to apply and obtain certification.

A. Duty to apply.  All owners [ or operators ] of noncommercial environmental laboratories shall apply for certification as specified by the provisions of this section.

B. Timely initial applications.

1. Owners [ or operators ] of noncommercial environmental laboratories applying for certification under this chapter for the first time shall submit an application to DGS-DCLS no later than [ 240 calendar days after the effective date of this chapter May 29, 2009 ].

2. Owners [ or operators ] of noncommercial environmental laboratories that come into existence after [ this chapter becomes effective October 1, 2008, ] shall submit an initial application to DGS-DCLS no later than 180 calendar days prior to beginning operation.

C. Timely renewal applications.  The owner [ or operator ] of [ an a ] certified noncommercial environmental laboratory shall submit an application for renewal of certification at least 90 calendar days prior to expiration of certification.

D. Responsibilities of the owner [ or and ] operator [ when the laboratory is owned by one person and operated by another person ].

1. When an environmental laboratory is owned by one person but is operated by another person, the operator may submit the application for the owner.

2. If the operator fails to submit the application, the owner is not relieved of his responsibility to apply for certification.

3. While DGS-DCLS may notify noncommercial environmental laboratories of the date their applications are due, failure of DGS-DCLS to notify does not relieve the owner [ or operator ] of his obligation to apply under this chapter.

E. Submission of applications for modifications to certification.  An owner [ or operator ] of a certified noncommercial environmental laboratory shall follow the process set out in 1VAC30-45-90 B to add a new [ technology matrix, technology/method ], an analyte or [ a test method analyte group ], modify a test method or institute use of a method [ or technology ] not in the laboratory’s standard operating procedures, including alternative test methods or procedures.

F. Contents of application.

1. Applications shall include the following information and documents:

a. Legal name of laboratory;

b. Name of owner of laboratory;

c. Name of operator of laboratory, if different than owner;

d. Street address and description of location of laboratory;

e. Mailing address of laboratory, if different from street address;

f. Address of owner, if different from laboratory address;

g. Name, address, telephone number, facsimile number and e-mail, as applicable, of responsible official;

h. Name, address, telephone number, facsimile number and e-mail, as applicable, of laboratory manager;

i. Name, address, telephone number, facsimile number and e-mail, as applicable, of designated quality assurance officer;

j. Name [ title, ] and telephone number of laboratory contact person;

k. Laboratory type (e.g., public water system, public wastewater system [  or combination of the two ] , or industrial (with type of industry indicated));

l. Laboratory hours of operation;

m. Fields of testing (program, test methods, and analytes) certification (matrix, technology/method, and analyte/analyte group) ] for which certification is sought;

n. Methods employed, including analytes;

o. The results of the three most recent proficiency test studies;

p. Quality assurance manual;

q. Lab identification number (for renewal only); and

r. For mobile laboratories, a unique vehicle identification number, such as a manufacturer’s vehicle identification number (VIN#), serial number, or license number.

2. Fee.  The application shall include payment of the fee as specified in 1VAC30-45-130.

3. Certification of compliance.

a. The application shall include a "Certification of Compliance" statement signed and dated by the responsible official, by the quality control officer and by the laboratory manager.

b. The certification of compliance shall state:  “The applicant understands and acknowledges that the laboratory is required to be continually in compliance with the Virginia environmental laboratory certification program regulation (1VAC30, Chapter 45) and is subject to the provisions of 1VAC30-45-100 in the event of noncompliance.  I certify under penalty of law that this document and all attachments were prepared under my direction or supervision in accordance with a system designed to assure that qualified personnel properly gather and evaluate the information submitted.  Based on my inquiry of the person or persons who manage the laboratory or those persons directly responsible for gathering and evaluating the information, the information submitted is, to the best of my knowledge and belief, true, accurate and complete.  Submitting false information or data shall result in denial of certification or decertification.  I hereby further certify that I am authorized to sign this application.”

G. Completeness determination.

1. DGS-DCLS shall determine whether an application is complete and notify the laboratory of the result of such determination. [ Except during During the initial certification period, DGS-DCLS shall provide this notice within 90 calendar days of its receipt of a laboratory's initial application. Following ] the initial certification period, DGS-DCLS shall provide this notice within 60 calendar days of DGS-DCLS’s receipt of [ the a laboratory's initial ] application [ and within 30 calendar days of DGS-DCLS' receipt of a laboratory's renewal application ].

2. An application shall be determined complete if it contains all the information required pursuant to subsection F of this section and is sufficient to evaluate the laboratory prior to the on-site assessment.  Designating an application complete does not preclude DGS-DCLS from requesting or accepting additional information.

3. If DGS-DCLS determines that an application is incomplete, DGS-DCLS's notification of such determination shall explain why the application is incomplete and specify the additional information needed to make the application complete.

4. Except during the initial certification period, if no determination is made within 60 calendar days of DGS-DCLS's receipt of either (i) the application or (ii) additional information, in the case of an application determined to be incomplete, the application shall be determined to be complete. [ During the initial certification period, the time period shall be 90 calendar days. ]

5. [ If the laboratory has not submitted the required additional information within 90 days of receiving a notice from DGS-DCLS requesting additional information, ] DGS-DCLS may [ deny any application from a laboratory and require the laboratory to submit a new application if the laboratory does not submit additional information required by DGS-DCLS within 90 days of receiving a notice that requires additional information return the incomplete application and inform the laboratory that the application cannot be processed. The laboratory may then submit a new application ].

H. Grant of interim certification pending final determination on application.

1. DGS-DCLS shall grant a laboratory interim certification status under the following conditions:

a. The laboratory’s application is determined to be complete;

b. The laboratory has satisfied all the requirements for certification, including all requests for additional information, with the exception of on-site assessment; and

c. DGS-DCLS is unable to schedule the on-site assessment within 90 days of its determination that the application is complete [ and that the laboratory has satisfied all other requirements for certification (for initial applications) or before the laboratory's certification expires (for renewal applications) ].

2. A laboratory with interim certification [ status ] shall have the same rights and status as a laboratory that has been granted certification by DGS-DCLS.

3. Interim certification expires when DGS-DCLS issues a final determination on certification.

I. On-site assessment.

[ 1. ] An on-site assessment shall be performed and the follow-up and reporting procedures for such assessments shall be completed in accordance with Article 2 (1VAC30-45-300 et seq.) of Part II of this chapter prior to issuance of a final determination on certification.

[ 2. Alternative on-site assessment option. If DGS-DCLS is unable to schedule an on-site assessment under the conditions of subsection H 1 c of this section, the owner of the applicant laboratory may use third-party on-site assessors instead of DGS-DCLS on-site assessors under the following conditions:

a. The third-party on-site assessors are on a DGS-DCLS-approved list of on-site assessors, and

b. The owner of the applicant laboratory agrees to pay the third-party on-site assessors. ]

J. Final determination on certification.

1. Upon completion of the certification review process and corrective action, if any, DGS-DCLS shall grant certification in accordance with subsection K of this section or deny certification in accordance with subsection L of this section.

2. Except during the initial certification period, DGS-DCLS shall complete action on a laboratory’s application within nine months from the time [ an a completed ] application is [ determined to be complete received from the laboratory ].

K. Grant of certification.

1. When a laboratory meets the requirements specified for receiving certification, DGS-DCLS shall issue a certificate to the laboratory.  The certificate shall be sent to the laboratory manager, and the responsible official shall be notified.

2. The [ director of DGS-DCLS shall sign the ] certificate [ shall be signed by the director of DGS-DCLS and. The certificate ] shall include the following information:

a. Name of owner [ or operator  ] of laboratory;

[ b. Name of ] operator of laboratory [ , if different from owner ];

[ b. c. ] Name of responsible official;

[ c. d. ] Address and location of laboratory;

[ d. e. ] Laboratory identification number;

[ e. f. ] Fields of [ testing (program, method certification (matrix, technology/method ] , [ analyte or other parameter analyte/analyte group ]) for which certification is granted;

f. Any addenda or attachments; and

g. Issuance date and expiration date.

3. The laboratory shall post the most recent certificate of certification and any addenda to the certificate issued by DGS-DCLS in a prominent place in the laboratory facility.

4. Certification shall expire two years after the date on which certification is granted.

L. Denial of certification.

1. DGS-DCLS shall deny certification to an environmental laboratory in total if the laboratory [ owner or an employee falsifies is found to be falsifying ] any data or [ provides providing ] false information to support certification.

2. Denial of certification in total or in part.

a. DGS-DCLS may deny certification to an environmental laboratory in total or in part if the laboratory [ owner or an employee ] fails to do any of the following:

(1) Pay the required fees.

(2) Employ laboratory staff to meet the personnel qualifications as required by Part II (1VAC30-45-200 et seq.) of this chapter.

(3) Successfully analyze and report proficiency testing samples as required by Part II of this chapter.

(4) Submit a corrective action report in accordance with Part II of this chapter in response to a deficiency report from the on-site assessment team within the required 30 calendar days.

(5) Implement the corrective actions detailed in the corrective action report within the time frame specified by DGS-DCLS.

(6) Pass required on-site assessment as specified in Part II of this chapter.

(7) Implement a quality system as defined in Part II of this chapter.

b. DGS-DCLS may deny certification to an environmental laboratory in total or in part if the laboratory’s application is not determined to be complete within 90 calendar days following notification of incompleteness because the laboratory is delinquent in submitting information required by DGS-DCLS in accordance with this chapter.

c. DGS-DCLS may deny certification to an environmental laboratory in total or in part if the DGS-DCLS on-site assessment team is unable to carry out the on-site assessment pursuant to Article 2 (1VAC30-45-300 et seq.) of Part II of this chapter because [ an employee, owner, or other a ] representative of the environmental laboratory denied the team entry during [ the laboratory's ] normal business hours [ that it specified in its application ].

3. [ To deny certification, DGS-DCLS shall provide by certified mail written notification of denial to the responsible official and manager of the laboratory, including a detailed explanation of the reason for denial and notice of the right to appeal such denial DGS-DCLS shall follow the process specified in 1VAC30-45-110 when denying certification to an environmental laboratory ].

M. Reapplication following denial of certification.

1. Upon denial of certification, the laboratory shall wait six months before reapplying for certification.

2. DGS-DCLS shall not waive application fees for a laboratory reapplying for certification.

1VAC30-45-80. Maintaining certification.

A. Certification remains in effect until withdrawn by DGS-DCLS, withdrawn voluntarily at the written request of the certified laboratory, or [ until ] expiration of the certification period.  To maintain certification, the certified laboratory shall comply with the elements listed in this section and in 1VAC30-45-90.

B. Quality systems.  Laboratories seeking to maintain certification under this chapter shall assure consistency and promote the use of quality assurance and quality control procedures.  Article 4 [ (1VAC30-600 1VAC30-45-600 ] et seq.) of Part II of this chapter specifies the quality assurance and quality control requirements that shall be met to maintain certification.

C. Proficiency tests.  Laboratories seeking to maintain certification under this chapter shall perform proficiency tests as required under Article 3 (1VAC30-45-500 et seq.) of Part II of this chapter.

D. Recordkeeping and retention.  All laboratory records associated with certification parameters shall be kept as provided by the requirements for records under Part II (1VAC30-45-200 et seq.) of this chapter.  These records shall be maintained for a minimum of three years unless the records are required to be maintained for a longer period by another section of this regulation or another regulation.  All such records shall be available to DGS-DCLS upon request.

1VAC30-45-90. [ Changing Notifications and changes to ] certification [ elements and ] status.

A. Changes to key certification criteria. [ 1. ] The certified laboratory shall notify DGS-DCLS [ as set out in subdivision 2 of this subsection in writing ] of any changes in key certification criteria within 30 calendar days of the change.  Key certification criteria are laboratory ownership, location, key personnel, [ test methods, analytes, ] and major instrumentation.

[ 2. The laboratory may initially notify DGS-DCLS of any change to key certification criteria by e-mail, facsimile or telephone.  The notification by e-mail, facsimile or telephone subsequently shall be submitted in writing.

3. As specified in subsection B of this section, changes to key certification criteria that affect the laboratory’s scope of certification require review and approval by DGS-DCLS in advance of the laboratory’s making the change. ]

B. Changes to scope of certification.

1. DGS-DCLS [ shall review and may ] approve [ the addition of a laboratory's application to add ] a new [ matrix, ] technology, [ an ] analyte, or [ a ] test method to a laboratory’s scope of certification [ or to otherwise modify the laboratory's scope of certification by performing a data review ] .

2. To [ begin the process of review apply ], the owner [ or operator ] of the certified laboratory [ that wants to add to the laboratory’s scope of certification ] shall submit the following [ application materials ] [ to DGS-DCLS:

a. A letter signed by the owner [ or operator ] that briefly summarizes the addition to be made to the laboratory’s scope of certification.

b. Pertinent information demonstrating [ that the laboratory is capable of performing the test method or using the technology to be added such as proficiency testing performance and quality control performance the laboratory's capability to perform the additional matrix, technology/method, or analyte/analyte group, such as proficiency testing performance and quality control performance ] .

c. A written standard operating procedure covering the new [ method, analyte, or technology matrix, technology/method, or analyte/analyte group ].

[ DGS-DCLS may request additional material to complete its review. ]

3. DGS-DCLS may approve a laboratory’s application for modification to its scope of certification by performing a review of the application materials submitted, without an on-site assessment. [ An The ] addition of a [ new ] technology or test method requiring [ the use of ] specific equipment may require an on-site assessment.  Other reviews of performance and documentation may be carried out by DGS-DCLS [ , ] depending on the modification for which the laboratory applies.

4. [ Within 90 calendar days of the receipt of the application from the certified environmental laboratory, DGS-DCLS shall review and determine whether the proposed modification may be approved.

5. ] If the proposed modification to the laboratory’s scope of certification is approved, DGS-DCLS shall amend the laboratory’s certificate of certification.

C. Change of ownership or location of laboratory.

1. The certified laboratory shall submit a written notification to DGS-DCLS of the change of ownership or location of the laboratory within 30 calendar days of the change. [ This requirement applies only to fixed-based and not mobile laboratories. ]

2. Certification may be transferred when the legal status or ownership of a certified laboratory changes [ without affecting its as long as the transfer does not affect the laboratory's ] personnel, equipment, [ and facilities or organization ].

3. [ DGS-DCLS may charge a transfer fee and may conduct an on-site assessment to verify the effects of such changes on laboratory performance If the laboratory's personnel, equipment, or organization are affected by the change of legal status or ownership, DGS-DCLS may require recertification or reapplication in any or all of the categories for which the laboratory is certified ].

[ 4. DGS-DCLS may require an on-site assessment depending on the nature of the change of legal status or ownership. DGS-DCLS shall determine the elements of any on-site assessment required. ]

[ 4. 5. ] When [ a laboratory changes there is a change in ] ownership, the new owner [ of the certified laboratory ] shall assure [ that the history of the laboratory ownership can be traced through historical traceability of the laboratory identification numbers. ]

[ 5. 6. When there is a change in ownership,  the new owner of the certified laboratory shall keep ] all records and analyses performed by the previous owner under his scope of certification [ shall be kept ] for a period of [ five three ] years [ . As required under 1VAC30-45-80 D, all such records shall be made available to DGS-DCLS upon request , or longer if required by other regulations. These records and analyses are subject to inspection by DGS-DCLS during this three-year period. This provision applies regardless of change of ownership, accountability or liability ].

D. Voluntary withdrawal.  Any environmental laboratory owner [ or operator ] who wishes to withdraw the laboratory from its certification status or from being certified, in total or in part, shall submit written notification to DGS-DCLS no later than 30 calendar days before the end of the laboratory’s certification term.  Within 30 calendar days, DGS-DCLS shall provide the laboratory with a written notice of withdrawal.

1VAC30-45-100. Decertification.

A. DGS-DCLS shall decertify an environmental laboratory in total [ for any of the following reasons: if the laboratory is found to be falsifying any data or providing false information to support certification. ]

[ 1. Submittal by the laboratory owner or an employee of proficiency test sample results generated by another laboratory as its own.

2. Falsification by a laboratory owner or an employee of any data or the provision of false information by any laboratory owner or an employee to support certification.

3. Conviction of the laboratory owner or an employee of charges relating to the falsification of any report concerning a laboratory analysis. ]

B. DGS-DCLS may decertify an environmental laboratory in part or in total when the laboratory [ owner or an employee ] has failed to do any of the following:

1. Participate in the proficiency testing program as required by Article 3 [ (1VAC30-40-500 1VAC30-45-500 ]  et seq.) of Part II of this chapter.

2. Complete proficiency testing studies and maintain a history of at least two successful proficiency testing studies for each affected certified field of testing out of the three most recent proficiency testing studies as defined in Article 3 (1VAC30-45-500 et seq.) of Part II of this chapter.

3. Maintain a quality system as defined in Article 4 (1VAC30-45-600 et seq.) of Part II of this chapter.

4. Employ staff that meet the personnel qualifications in Article 1 (1VAC30-45-200 et seq.) of Part II of this chapter.

5. Submit an acceptable corrective action report after two opportunities as specified in 1VAC30-45-390.

6. Implement corrective action specified in the laboratory’s corrective action report as set out under 1VAC30-45-390.

7. Notify DGS-DCLS of any changes in key certification criteria as set forth in 1VAC30-45-90.

8. Use accurate references to the laboratory’s certification status in the laboratory’s documentation.

C. [ To decertify an environmental laboratory, DGS-DCLS shall provide by certified mail written notification of the decertification to the responsible official and manager of the laboratory, including a detailed explanation of the reason for the decertification and notice of the right to appeal such decertification DGS-DCLS shall follow the process specified in 1VAC30-45-110 when decertifying an environmental laboratory ].

D. Responsibilities of the environmental laboratory and DGS-DCLS when certification has been withdrawn.

1. Laboratories that lose their certification in full shall return their certificate to DGS-DCLS.

2. If a laboratory loses certification in part, an addendum to the certificate shall be issued by DGS-DCLS to the laboratory.

E. After correcting the reason or cause for decertification under 1VAC30-45-100 A or B, the laboratory owner [ or operator ] may reapply for certification.

1VAC30-45-110. [ Appeal procedures Procedures to deny certification, to decertify a laboratory, and appeal procedures ].

A. [ If DGS-DCLS believes it has grounds to deny certification or to decertify an environmental laboratory, ] DGS-DCLS shall notify [ an the ] environmental laboratory in writing of its [ decision to deny certification or to decertify an environmental laboratory intent to hold an informal fact finding under §2.2-4019 of the Code of Virginia in order to make a decision on the denial of certification or decertification. DGS-DCLS shall send this notification by certified mail to the responsible official and provide a copy to the manager of the environmental laboratory. The notice of informal fact finding shall provide a detailed explanation of the basis for the notice ] .

B. [ All appeals taken from actions of the DGS-DCLS director relative to the provisions of this chapter shall be governed by the Virginia Administrative Process Act (§2.2-4000 et seq. of the Code of Virginia) . Following the informal fact finding held pursuant to §2.2-4019 of the Code of Virginia, the director shall render a decision regarding certification, and shall send this notification by certified mail to the responsible official and provide a copy to the manager of the environmental laboratory. If the director’s decision is adverse to the environmental laboratory, the responsible official may appeal this decision in accordance with §2.2-4026 of the Code of Virginia and Part 2A of the Rules of the Supreme Court of Virginia.

C. The provisions of this section do not preclude informal discussions between DGS-DCLS and any environmental laboratory that has been notified of a possible denial of certification or of decertification. These informal discussions to resolve the concerns that prompted the notice shall be held prior to the informal fact-finding proceeding.

D. The certification status of an environmental laboratory appealing decertification shall not change pending the final decision of the appeals filed under the Virginia Administrative Process Act (§2.2-4000 et seq. of the Code of Virginia) and Part 2A of the Rules of Supreme Court of Virginia. ]

1VAC30-45-120. Exemptions.

A. DGS-DCLS may grant a partial or full exemption from the requirements of this chapter based on compliance and performance.

B. DGS-DCLS may consider granting an exemption if a laboratory applies for an exemption and has met all certification requirements for a period of four consecutive years.

C. An environmental laboratory may apply for an exemption by submitting a request.  The request shall include the following information:

1. The scope of the requested exemption;

2. Whether the exemption should be partial or total;

3. If partial, what form the exemption will take; and

4. Why the exemption is appropriate.

D. Upon receiving an application for an exemption, DGS-DCLS shall provide notice of the request for an exemption in the Virginia Register of Regulations.

E. The notice shall provide a 30-day comment period on the request and shall specify the nature of the request.

F. DGS-DCLS shall grant or deny the exemption request and provide a written response to the requesting laboratory within 90 calendar days of receipt of the request.

G. Exemptions granted by DGS-DCLS shall be for a period of no more than 24 months.

1VAC30-45-130. Fees.

A. General.

1. Fees shall be submitted with all applications [ , including reapplications, ] for certification  [ and all renewal applications for certification ].  Applications shall not be designated as complete until the fee is received by DGS-DCLS.

2. Fees shall be nonrefundable.

B. Fee computation.

1. Fees shall be computed based on the test methods for which a laboratory seeks certification and on the laboratory type. For the purpose of fee calculation, the designations for the laboratory type are (i) a general environmental laboratory or (ii) an environmental laboratory performing only simple test procedures.

2. The fee shall be the total of the base fee and the test category fees for the specific laboratory type to be certified.

3. The test category fees cover categories for the test methods to be certified as specified in the laboratory’s application.

4. If the total of the base fee and the test category fees is more than the maximum fee designated for the specific laboratory type to be certified, the laboratory shall pay the maximum fee.

C. Laboratories performing only simple test procedures.

1. The base fee shall be $100.

2. The maximum fee shall be [ $400 $600 ].

D. General environmental laboratories.

1. The base fee shall be $1,700.

2. The maximum fee shall be [ $3,800 $5,200 ].

E. Test category fees.

1. Fees shall be charged for each category of tests to be certified.

2. The fee for each category includes one or more analytical methods unless otherwise specified.  With the exception of the test categories labeled oxygen demand and physical, test categories related to test methods for water are defined by 40 CFR 136.3.

3. Fees.

| TEST CATEGORY | FEE |
| --- | --- |
| Oxygen demand (BOD or COD) | [ $300 $375 ] |
| Bacteriology | [ $300 $375 ] |
| Inorganic chemistry, fewer than four methods | [ $300 $375 ] |
| Inorganic chemistry, four or more methods | [ $600 $750 ] |
| Chemistry metals, [ fewer than four one - two ] methods | [ $300 $450 ] |
| Chemistry metals, [ four or ] more [ than two ] methods | [ $600 1,000 ] |
| Organic chemistry, fewer than four methods | [ $350 $600 ] |
| Organic chemistry, four or more [ methods ] | [ $700 $1,200 |
| [ Whole effluent Aquatic ] toxicity, acute methods only | [ $300 $400 ] |
| [ Whole effluent Aquatic ] toxicity, acute and chronic methods | [ $600 $700 ] |
| Radiochemical | [ $900 $1,000 ] |
| Physical | [ $300 $375 ] |

F. Additional fees. [ Additional fees shall be charged to laboratories applying for the following: (i) modification to scope of certification under 1VAC30-45-90 B, (ii) transfer of ownership under 1VAC30-45-90 C, (iii) exemption under 1VAC30-45-120, (iv) request that multiple noncontiguous laboratory sites be considered as one site under 1VAC30-45-60 B 3, or (v) petition for a variance under 1VAC30-45-140. ]

1. General environmental laboratories applying for an exemption under 1VAC30-45-120 shall pay an application fee of $250 and if the exemption is granted, up to an additional $1,000 depending on the scope of the exemption.  Laboratories performing only simple test procedures applying for an exemption under 1VAC30-45-120 shall pay an application fee of $100 and if the exemption is granted, up to an additional $1,000 depending on the scope of the exemption. The fee assessed for the scope of the exemption shall be based on the actual time needed for DGS-DCLS to make the determination. The fee assessed shall be calculated using the method in [ subdivision 4 subsection G ] of this [ subsection section ].

2. For any certified environmental laboratory that applies to modify its scope of certification as specified under 1VAC30-45-90 B, DGS-DCLS shall assess a fee determined by the method in [ subdivision 4 subsection G ] of this [ subsection section ].

3. Under 1VAC30-45-90 C, DGS-DCLS may charge a transfer fee to a certified laboratory that transfers ownership. [ A fee shall be charged if DGS-DCLS (i) needs to review documentation sent by the laboratory about the transfer of ownership or (ii) determines that an on-site assessment is necessary to evaluate the effect of the transfer of ownership. DGS-DCLS shall assess a fee determined by the method in subsection G of this section. ] If DGS-DCLS determines that a fee should be charged, the fee shall be a minimum of $100 and a maximum of $1,000. [ If DGS-DCLS determines that an on-site assessment is necessary to evaluate the effect of the transfer of ownership, DGS-DCLS shall assess a fee determined by the method in subdivision 4 of this subsection. If, under 1VAC30-45-90 C, DGS-DCLS determines that the change of ownership or location of laboratory requires recertification of or reapplication by the laboratory, the laboratory shall pay the application fees required under this section.

4. Under 1VAC30-45-60 B 3, the owner of multiple noncontiguous laboratories may request that DGS-DCLS consider these laboratories to be one site. If, as a result of the request being granted, DGS-DCLS needs to perform multiple on-site assessments, DGS-DCLS shall charge a fee for the additional on-site assessments. The fee shall be the sum of reasonable travel costs and labor charges for the additional on-site assessments. The labor charges will be determined following the method in subsection G of this section.

5. Under 1VAC30-45-140, any person regulated by this chapter may petition the director to grant a variance from any requirement of this chapter. DGS-DCLS shall charge a fee for the time needed to review the petition, including any on-site assessment required. The fee shall be determined by the method specified in subsection G of this section. ]

[ 4. G. ] Fee determination.

[ a. 1. ] The fee shall be the sum of the total hourly charges for all reviewers plus any on-site review costs incurred.

[ b. 2. ] An hourly charge per reviewer shall be determined by (i) obtaining a yearly cost by multiplying the reviewer’s annual salary by 1.35 (accounts for overhead such as taxes and insurance) and then (ii) dividing the yearly cost by 1,642 (number of annual hours established by Fiscal Services, DGS, for billing purposes).

[ c. 3. ] The charge per reviewer shall be determined by multiplying the number of hours expended in the review by the reviewer’s hourly charge.

[ d. 4. ] If an on-site review is required, travel time and on-site review time shall be charged at the same hourly charge per reviewer, and any travel expenses shall be added.

[ G. On-site assessment fees.  When, with the concurrence of the applicant laboratory, DGS-DCLS uses approved, third-party on-site assessors, the cost of the on-site assessment shall be paid by the applicant. H. Out-of-state laboratories - travel costs. The owner of an environmental laboratory located in another state who applies for certification under this chapter shall also pay a fee equal to the reasonable travel costs associated with conducting an on-site assessment at the laboratory. Reasonable travel costs include transportation, lodging, per diem, and telephone and duplication charges.

I. DGS-DCLS shall derive the travel costs charged under subsections G and H of this section from the Commonwealth of Virginia reimbursement allowances and rates for lodging, per diem, and mileage. ]

1VAC30-45-140. Petitioning for a variance.

A. Any person regulated by this chapter may petition the director to grant a variance from any requirement of this chapter. Any person submitting a petition to the director [ must shall ] meet the provisions of this section.  Any petition submitted to the director is subject to the Virginia Administrative Process Act (§2.2-4000 et seq. of the Code of Virginia).

B. The petition shall be submitted to the director by certified mail and shall include:

1. The petitioner's name and address;

2. A statement of the petitioner’s interest in the proposed action;

3. A description of desired action and a citation of the regulation from which a variance is requested;

4. A description of need and justification for the proposed action, including impact of the proposed action on the laboratory’s operation;

5. Information demonstrating that the requested variance will meet the purposes and objectives of the relevant regulatory provision and of §2.2-1105 of the Code of Virginia (Environmental Laboratory Certification Program);

6. The duration of the variance, if applicable;

7. The potential impact of the variance on public health or the environment;

8. Other information believed by the applicant to be pertinent; and

9. The following statement signed by the petitioner or authorized representative:  "I certify that I have personally examined and am familiar with the information submitted in this petition and all attached documents, and that, based on my inquiry of those individuals immediately responsible for obtaining the information, I believe that the submitted  information is true, accurate, and complete. I am aware that there are significant penalties for submitting false information, including the possibility of fine and imprisonment."

C. Petition processing.

1. After receiving a petition that includes the information required in subsection B of this section, the director will determine whether the information received is sufficient to render the decision.  If the information is deemed insufficient, the director will specify additional information needed and request that it be furnished.

2. The petitioner may submit the additional information requested, or may attempt to show that no reasonable basis exists for additional information. If the director agrees that no reasonable basis exists for the request for additional information, he will act in accordance with subsection D of this section. If the director [ continues to believe finds ] that a reasonable basis exists to require the submission of such information, he will proceed with the denial action in accordance with the Administrative Process Act.

D. Public review of tentative decision. The director will evaluate the application and issue a draft notice tentatively denying the petition, granting the variance as requested, or granting a modified or partial variance. Notification of this tentative decision will be published in the Virginia Register of Regulations.  The director will accept comment on the tentative decision for 30 days, and shall hold a public hearing if a request is received or at his discretion if there is no request. The director will issue a final decision after receipt of comments and after the hearing (if any).

E. Conditions for granting variance request or a modified variance.

1. The director may grant the variance if the applicant demonstrates to the satisfaction of the director that:

a. The proposed variance will meet the goals and purposes of the provisions from which a variance is sought; and

b. The variance does not conflict with federal or state law or regulations.

2. If the director grants a variance request, the notice to the petitioner shall provide that the variance may be terminated upon a finding by the director that the petitioner has failed to comply with any requirements of the variance.

3. When a modified variance is granted, the director may:

a. Specify the termination date of the variance;

b. Include a schedule for:

(1) Compliance, including increments of progress, by the laboratory with each requirement of the variance; and

(2) Implementation by the laboratory of such measures as the director finds necessary in order that the variance may be granted.

F. Decisions to grant or deny a petition [ in whole or in part, or to modify or terminate a variance ] are subject to the provisions of Article 3 (§2.2-4018 et seq.) of the Virginia Administrative Process Act.

1VAC30-45-150. (Reserved.)

1VAC30-45-160. (Reserved.)

1VAC30-45-170. (Reserved.)

1VAC30-45-180. (Reserved.)

1VAC30-45-190. (Reserved.)

Part II 
Standards

Article 1 
Personnel

1VAC30-45-200. Laboratory manager.

A. Laboratory manager - general.

1. Each environmental laboratory shall designate a person to be responsible for the general oversight of the operation of the laboratory in accordance with this chapter, including the day-to-day functioning and administration of the laboratory, the technical operations, supervision of laboratory procedures, reporting of laboratory results, and implementation of any corrective actions.

2. The title of this person may include but is not limited to laboratory director, technical director, laboratory supervisor or laboratory manager.

[ 3. A laboratory may appoint one or more technical directors for the appropriate fields of certification for which the laboratory seeks certification. ]

B. Laboratory manager - qualifications.

1. For an environmental laboratory that performs procedures beyond simple test procedures, a laboratory manager shall have two years of experience managing an environmental laboratory or performing the analyses for which the environmental laboratory seeks certification or both.

2. For an environmental laboratory that performs only simple test procedures, [ there are no qualification requirements for ] a laboratory manager [ except that the responsible official ] shall [ be designated by the responsible official designate the laboratory manager ].

[ 3. A full-time employee of a drinking water or wastewater treatment facility who holds a valid treatment plant operator’s certificate appropriate to the nature and size of such facility shall be deemed to meet the educational and experience requirements serving as the laboratory manager of the certified laboratory devoted exclusively to the examination of environmental samples taken within such facility system and limited to the scope of that facility’s regulatory permit.

4. A full-time employee of an industrial waste treatment facility with a minimum of one year of experience under supervision in environmental analysis shall be deemed to meet the requirements for serving as the laboratory manager of a certified laboratory devoted exclusively to the examination of environmental samples taken within such facility for the scope of that facility’s regulatory permit. ]

1VAC30-45-210. Quality assurance officer.

A. The laboratory shall have a quality assurance officer who shall be responsible for the quality system and its implementation.  Where staffing is limited, the quality assurance officer may also be the laboratory manager.  The quality assurance officer may be employed on a part-time basis or be a consultant.

B. The quality assurance officer shall have documented training or experience in quality assurance and quality control procedures and be knowledgeable in the quality system as defined in Article 4 (1VAC30-45-600 et seq.) of Part II of this chapter.  The quality assurance officer shall have a general knowledge of the analytical test methods for which data review is performed.

C. The responsibilities of the quality assurance officer shall include, but not be limited to, the implementation and oversight of the quality system, the implementation of new quality assurance and control practices, periodic audits of the quality system in place, periodic review of final data reports, and documentation of laboratory quality system deficiencies.

1VAC30-45-220. Laboratory personnel requirements and management responsibilities.

A. The laboratory shall have sufficient personnel, having the necessary education, training, technical knowledge and experience for their assigned functions.

B. The laboratory manager shall ensure that the training of the laboratory personnel is kept up to date.

C. Laboratory personnel shall be responsible for complying with all quality systems requirements set out in Article 4 (1VAC30-45-600 et seq.) of Part II of this chapter that are pertinent to their assigned functions.

D. The laboratory manager shall ensure that laboratory personnel have demonstrated initial and ongoing capability to perform their assigned functions. [ See 1VAC30-45-730 E and F. ]

E. Records on the relevant qualifications, training skills and experience of the laboratory personnel, including records on demonstrated proficiency for each test method, shall be maintained by the laboratory manager.

1VAC30-45-230. Absence of laboratory manager [ or quality assurance officer ].

[ The When a laboratory manager will be absent for a period exceeding 15 consecutive calendar days, the ] laboratory shall [ nominate deputies in the case of absence of the laboratory manager or the quality assurance officer designate a qualified replacement to perform the manager's function ].

1VAC30-45-240. (Reserved.)

1VAC30-45-250. (Reserved.)

1VAC30-45-260. (Reserved.)

1VAC30-45-270. (Reserved.)

1VAC30-45-280. (Reserved.)

1VAC30-45-290. (Reserved.)

Article 2 
On-Site Assessment

1VAC30-45-300. Frequency of on-site assessment.

A. A comprehensive on-site assessment shall be conducted of each laboratory as a condition for granting certification [ initially and at renewal every two years ].

B. [ On-site Other on-site ] assessments [ may be conducted more frequently for cause ].

1. [ Situations that might trigger more frequent on-site assessments include review of a previously deficient on-site assessment, poor performance on a proficiency testing sample, change in other certification elements, or other information concerning the capabilities or practices of the certified laboratory If DGS-DCLS identified a deficiency on a previous on-site assessment, the agency may conduct a follow-up on-site assessment ] .

[ 2. DGS-DCLS may reassess a laboratory prior to taking a regulatory or administrative action affecting the laboratory's certification.

3. An 2. DGS-DCLS may conduct an on-site ] assessment [ may be conducted ] when a [ major laboratory applies to modify its scope of certification, when a transfer of owner occurs that affects personnel, equipment, or the laboratory facilities, or when a laboratory applies for an exemption or a variance. Any other ] change [ occurs occurring ] in a laboratory's operations that might reasonably be expected to alter or impair analytical capability and quality [ may trigger an on-site assessment ].

1VAC30-45-310. Announced and unannounced on-site assessments.

A. DGS-DCLS may conduct, at its discretion, either announced or unannounced on-site assessments.

B. Advance notice of an assessment shall not be necessary.

C. To the maximum extent practical, DGS-DCLS, when necessary, shall work with the owner [ or operator ] of an environmental laboratory to obtain government security clearances for assessment personnel as far in advance as possible.  The owner [ or operator ] of the environmental laboratory shall facilitate expeditious attainment of the necessary clearances.

D. To the maximum extent practical, assessment personnel shall minimize disruption of a laboratory's operations and take into account competing demands on the time of laboratory personnel.

1VAC30-45-320. Request for records.

Prior to the actual site visit, DGS-DCLS may request in writing from a laboratory those records required to be maintained by this chapter.

1VAC30-45-330. Areas to be assessed.

[ A. DGS-DCLS shall assess the laboratory against the personnel and quality control standards in Article 1 (1VAC30-45-200 et seq.) and Article 4 (1VAC30-45-600 et seq.) of this part. ] The [ specific ] areas evaluated in an on-site assessment shall include [ but not be limited to ]:

1. Adequacy of the laboratory facility.

2. Organization and management of the laboratory.

3. Qualifications and experience of laboratory personnel.

4. Receipt, tracking and handling of samples.

5. Quantity, condition, and performance of laboratory instrumentation and equipment.

6. Preparation and traceability of calibration standards.

7. Test methods [ (Including (including ] the adequacy of the laboratory’s standard operating procedures as well as confirmation of the analyst’s adherence to SOPs, and the analyst’s proficiency with the described task).

8. Data reduction procedures, including an examination of raw data and confirmation that final reported results can be traced to the raw data/original observations.

9. Quality assurance and quality control procedures, including adherence to the laboratory's quality assurance plan and adequacy of the plan.

[ 10. Recordkeeping. ]

[ B. These areas shall be evaluated against the standards set out in Article 4 (1VAC30-45-600 et seq.) of Part II of this chapter and the appropriate reference methods. ]

1VAC30-45-340. National security considerations.

A. Assessments at facilities owned or operated by federal agencies or contractors may require security clearances, appropriate badging, or a security briefing before the assessment begins.

B. The laboratory shall notify DGS-DCLS in writing of any information that is controlled for national security reasons and cannot be released to the public.

1VAC30-45-350. Arrival, admittance and opening conference.

A. Arrival. Assessment personnel shall arrive at the laboratory during established working hours.  The laboratory [ supervisor manager ] (or, if unavailable, the laboratory [ supervisor's manager's ] designee) shall be located as soon as possible after the assessment personnel arrive on the premises.

B. Admittance of assessment personnel. A laboratory's refusal to admit the assessment personnel for an on-site assessment shall result in an automatic failure of the laboratory to receive certification or loss of an existing certification by the laboratory, unless there are extenuating circumstances that are accepted and documented by DGS-DCLS. The team leader for the assessment personnel shall notify DGS-DCLS as soon as possible after refusal of entry.

C. Health and safety.

1. Under no circumstance, and especially as a precondition to gain access to a laboratory, shall assessment personnel be required or even allowed to sign any waiver of responsibility on the part of the laboratory for injuries incurred during an assessment.

2. Assessment personnel shall comply with all facility and laboratory safety procedures.

D. Opening conference. An opening conference shall be conducted and shall address the following topics:

1. The purpose of the assessment;

2. The identification of assessment personnel;

3. The test methods that will be examined;

4. Any pertinent records and procedures to be examined during the assessment and the names of the individuals in the laboratory responsible for providing assessment personnel with such records;

5. The roles and responsibilities of laboratory staff and managers;

6. Any special safety procedures that the laboratory may think necessary for the protection of assessment personnel;

7. The standards and criteria that will be used in judging the adequacy of the laboratory operation;

8. Confirmation of the tentative time for the exit conference; and

9. Discussion of any questions the laboratory may have about the assessment process.

1VAC30-45-360. On-site laboratory records review and collection.

A. Records shall be reviewed by assessment personnel for accuracy, completeness and the use of proper methodology for each analyte and test method to be evaluated.

B. Records required to be maintained pursuant to this chapter shall be examined as part of an assessment for certification.

1VAC30-45-370. Observations of and interviews with laboratory personnel.

A. As an element of the assessment process, the assessment team shall evaluate an analysis regimen by requesting that the analyst normally conducting the procedure give a step-by-step description of exactly what is done and what equipment and supplies are needed to complete the regimen.  Any deficiencies shall be noted and discussed with the analyst.  In addition, the deficiencies shall be discussed in the closing conference.

B. Assessment personnel may conduct interviews with appropriate laboratory personnel.

C. Calculations, data transfers, calibration procedures, quality control and quality assurance practices, adherence to test methods, and report preparation shall be assessed for the complete scope of certification with appropriate laboratory analysts.

1VAC30-45-380. Closing conference.

A. Assessment personnel shall meet with representatives of the laboratory following the assessment for a closing conference.

B. During the closing conference, assessment personnel shall inform the laboratory of the preliminary findings and the basis for such findings.  The laboratory shall have an opportunity to provide further explanation or clarification relevant to the preliminary findings.  If the laboratory objects to the preliminary findings during the closing conference, all objections shall be documented by the assessment personnel and included in the final report to DGS-DCLS.

C. Additional problem areas may be identified in the final report.

D. Any potentially illegal activity that may be the subject of further action shall not be discussed in the closing conference.

1VAC30-45-390. Follow-up and reporting procedures.

A. DGS-DCLS shall present an assessment report to the laboratory within 30 calendar days of the assessment.

B. If there are deficiencies identified in the assessment report, the laboratory shall have 30 calendar days from the date of its receipt of the assessment report to provide a response to DGS-DCLS.  This response shall be called a corrective action report.

C. An exception to the deadlines specified in subsections A and B of this section may occur in appropriate circumstances.  Two circumstances that may be considered appropriate by DGS-DCLS are where a possible enforcement investigation or other action has been initiated or where the laboratory shows good cause for an extension.

D. The corrective action report shall include the following:

1. Any objections that the laboratory has with regard to the assessment report;

2. The action that the laboratory proposes to implement to correct each deficiency identified in the assessment report; and

3. The time period required to accomplish the corrective action.

E. DGS-DCLS shall determine and shall notify the laboratory within 30 calendar days of receipt whether the corrective action report is an acceptable response to the deficiencies identified in the assessment report.

F. If the corrective action report (or a portion of the report) is determined to be unacceptable to remedy the deficiency, DGS-DCLS shall provide written notification to the responsible official and [ technical director manager ] of the laboratory including a detailed explanation of the basis for such determination.  Following receipt of such notification, the laboratory shall have an additional 30 calendar days to submit a revised corrective action report acceptable to DGS-DCLS.

1VAC30-45-400. Documentation of on-site assessment.

A. Checklists.  The checklists used by assessment personnel during the assessment shall become a part of DGS-DCLS's file for the laboratory.

B. Assessment report format.

1. The final assessment report shall contain a narrative description of the adequacy of the laboratory as it relates to the assessment standards specified in this chapter and in 1VAC30-45-330.

2. Assessment reports shall contain:

a. Name of owner [ or operator ] of the laboratory [ (or operator of the laboratory, if different from the owner) ];

b. Identification of the laboratory assessed;

c. Date of the assessment;

d. Identification and affiliation of all assessment personnel;

e. Identification of participants in the assessment process;

f. Identification of analytes and test methods assessed;

g. Statement of the objective of the assessment;

h. Summary;

i. Assessment observations, findings (including any deficiencies), objections noted by the laboratory, and requirements; and

j. Comments and recommendations.

3. The assessment findings and requirements shall be referenced to the standards in Part II (1VAC30-45-200 et seq.) of this chapter so that both the finding is understood and the specific requirement is outlined.  The assessor shall specify the laboratory records, documents, equipment, procedures, or staff evaluated and the observations that contributed to each identified deficiency.  The assessment report shall support with sufficient data all assessment findings and the overall evaluation of the laboratory.

4. The comments and recommendations section may be used to convey recommendations aimed at helping the laboratory improve.

C. Release of report.

1. The assessment report shall be released [ initially ] by DGS-DCLS to the [ responsible official and the ] laboratory [ supervisor manager ].  The assessment report shall not be released to the public until findings of the assessment and the corrective actions have been finalized, all information relating to national security has been stricken from the report in accordance with prescribed procedures, and the report has been provided to the laboratory.

2. [ Any documentation determined to be relevant to an ongoing enforcement investigation shall be considered exempt from release to the public Once the assessment report has been released to the laboratory, any member of the public may request a copy of the report under the requirements of the Virginia Freedom of Information Act (§2.2-3700 et seq. of the Code of Virginia) ].

3. Checklists used by assessment personnel during the on-site assessment shall be provided to the laboratory with the final on-site assessment report.

D. The laboratory shall have access to documentation pertaining to any on-site assessment of its facilities.  Any laboratory wishing to review its files shall request such assistance of DGS-DCLS five days prior to visiting DGS-DCLS.  A laboratory may request copies of its documents without visiting DGS-DCLS.  A reasonable fee may be charged for copying, mailing, and staff time.

1VAC30-45-410. (Reserved.)

1VAC30-45-420. (Reserved.)

1VAC30-45-430. (Reserved.)

1VAC30-45-440. (Reserved.)

1VAC30-45-450. (Reserved.)

1VAC30-45-460. (Reserved.)

1VAC30-45-470. (Reserved.)

1VAC30-45-480. (Reserved.)

1VAC30-45-490. (Reserved.)

Article 3 
Proficiency Testing

1VAC30-45-500. Laboratory enrollment in proficiency testing program.

A. Required level of participation.

1. To be certified initially and to maintain certification, a laboratory shall participate in two single-blind, single-concentration PT studies, where available, per year for each PT field of testing for which it seeks or wants to maintain certification. [ Laboratories applying to be certified for environmental toxicology (aquatic toxicity, sediment toxicity, or soils toxicity) shall meet the requirements of subdivision 3 of this subsection. ]

2. Laboratories shall obtain PT samples from [ any PT provider approved under the requirements of the NELAC standards for proficiency test providers set out in Chapter 2 of the 2003 standards such as ] NIST [ or another provider approved by DGS-DCLS ]. [ For PT fields of testing having no approved providers listed by NELAC, the laboratory shall consult DGS-DCLS for an approved provider. ]

3. [ Each laboratory shall participate in at least two PT studies for each PT field of testing per year unless DGS-DCLS approves a different frequency for a given program. Laboratories applying to be certified for environmental toxicology (aquatic toxicity, sediment toxicity, or soils toxicity). To be certified initially and to maintain certification, a laboratory shall participate in at least one PT study per year (i.e., not more than 12 months apart), when available, for each method code (matrix, organism, exposure system and endpoint) for which it seeks or wants to maintain certification. Laboratories seeking certification for aquatic toxicity testing shall meet the requirements of 1VAC30-45-530. ]

B. Requesting certification.

1. [ At the time each laboratory applies When applying ] for certification, [ it the laboratory owner ] shall notify DGS-DCLS [ for which of the ] fields of testing [ it for which the laboratory ] chooses to become certified and shall participate in the appropriate PT studies.

2. For all fields of testing for which PT samples are not available, the laboratory shall ensure the reliability of its testing procedures by maintaining a quality system that meets all applicable requirements of Article 4 (1VAC30-45-600 et seq.) of Part II of this chapter.

C. Reporting results.

1. Each laboratory shall authorize the PT study provider to release all certification and remediation results and "acceptable" or "not acceptable" status directly to DGS-DCLS, in addition to the laboratory.

2. The results of all of the PT sample tests including "acceptable" or "not acceptable" status shall be part of the public record.

[ 3. The result of a PT sample, "acceptable" or "not acceptable" status, shall apply to all certified methods within that matrix that a laboratory employs for an analyte. ]

1VAC30-45-510. Requirements for laboratory testing of PT study samples.

A. The samples shall be analyzed and the results returned to the PT study provider no later than 45 calendar days from the scheduled study shipment date. [ Samples for environmental toxicology shall be analyzed within 45 calendar days of sample receipt. The laboratory shall report the result within 45 calendar days of completion of the PT. ]

B. The laboratory’s management and all analysts shall ensure that all PT samples are managed, analyzed, and reported in the same manner as real environmental samples utilizing the same staff, methods as used for routine analysis of that analyte, procedures, equipment, [ and ] facilities [ , and frequency of analysis. When analyzing a PT sample, the laboratory shall employ the same calibration, laboratory quality control and acceptance criteria, sequence of analytical steps, number of replicates and other procedures as used when analyzing routine samples. ]

C. Restrictions on exchanging information.  Laboratories shall comply with all of the following restrictions on the transfer of PT samples and communication of PT sample results prior to the time the results of the study are released.  Laboratory management or staff shall not:

1. Send any PT sample, or a portion of a PT sample, to another laboratory for any analysis for which it seeks certification or is certified.

2. Knowingly receive any PT sample or portion of a PT sample from another laboratory for any analysis for which the sending laboratory seeks certification or is certified.

3. Communicate with any individual at another laboratory (including intracompany communication) concerning the PT sample.

4. Attempt to obtain the assigned value of any PT sample from their PT provider.

D. Maintenance of records.  The laboratory shall maintain copies of all written, printed, and electronic records, including but not limited to bench sheets, instrument strip charts or printouts, data calculations, and data reports, resulting from the analysis of any PT sample for [ five three ] years or for as long as is required by the applicable regulatory program.  These records shall include a copy of the PT study report forms used by the laboratory to record PT results.  All of these laboratory records shall be made available to the assessors of DGS-DCLS during on-site audits of the laboratory.

1VAC30-45-520. PT criteria for laboratory certification.

A. Result categories.

1. The criteria described in this section apply individually to each PT field of testing, as defined by the laboratory seeking to obtain or maintain certification in its certification request. These criteria apply only to the PT portion of the overall certification standard.

2. There are two PT result categories: "acceptable" and "not acceptable."

B. Initial and continuing certification.

1. A laboratory seeking to obtain or maintain certification shall successfully complete two PT studies for each requested PT field of testing within the most recent three rounds attempted.

2. Once a laboratory has been granted certification status, it shall continue to complete PT studies for each PT field of testing and maintain a history of at least two acceptable PT studies for each PT field of testing out of the most recent three.

3. For a laboratory seeking to obtain initial certification, the most recent three rounds attempted shall have occurred within 18 months of the laboratory’s application date.

4. For a laboratory seeking initial certification, or for a laboratory performing supplemental testing, the PT studies shall be at least [ 30 15 ] calendar days apart [ from the closing date of one study to the shipment date of another study for the same PT field of testing ].

5. For a laboratory to maintain certification, completion dates of successive proficiency rounds for a given PT field of testing shall be approximately six months apart. Failure to meet the semiannual schedule is regarded as a failed study.

C. Supplemental studies.

1. A laboratory may elect to participate in PT studies more frequently than required by the semiannual schedule. This may be desirable, for example, when a laboratory first applies for certification or when a laboratory fails a study and wishes to quickly reestablish its history of successful performance.

2. These additional studies shall be reported and are counted and scored the same way as routinely scheduled studies and shall be at least [ 30 15 ] calendar days apart.

D. Failed studies and corrective action.

1. Whenever a laboratory fails a study, it shall determine the cause for the failure and take any necessary corrective action. It shall then document in its own records and provide to DGS-DCLS both the investigation and the action taken.

2. If a laboratory fails two out of the three most recent studies for a given field of testing, its performance is considered unacceptable for that field. The laboratory shall then meet the requirements of initial certification as described in subsection B of this section.

E. Second failed study.

1. The PT provider reports laboratory PT performance results to DGS-DCLS at the same time that it reports the results to the laboratory.

2. If a laboratory fails a second study out of the most recent three, as described in subdivision D 2 of this section, DGS-DCLS shall take action within 60 calendar days to determine the certification status [ of all methods ] for the unacceptable [ analyte or analytes for that program and matrix PT field of testing ].

F. Scheduling of PT studies.

[ 1. DGS-DCLS may specify which months that laboratories within its authority are required to participate in PT study programs.

2. If DGS-DCLS chooses to specify the months, then it shall adhere to the required semiannual schedule. If DGS-DCLS does not specify the months, then the laboratory Laboratories shall determine the schedule for their PT studies ].

G. Withdrawal from PT studies. A laboratory may withdraw from a PT study for an analyte or analytes or for the entire study if the laboratory notifies both the PT provider and DGS-DCLS before the closing date of the PT study. This does not exempt the laboratory from participating in the semiannual schedule.

1VAC30-45-530 [ (Reserved.) Special requirements for aquatic toxicity.

A. Laboratories seeking certification for aquatic toxicity testing shall be assessed through on-site assessment and evaluation of EPA Discharge Monitoring Report-Quality Assurance (DMR-QA) test results when available. A failed DMR-QA endpoint shall require both of the following:

1. A formal response to DGS-DCLS with an explanation of the probable cause for the endpoint failure and description of corrective actions to be taken (where appropriate).

2. A decision by DGS-DCLS to accept the response or require additional actions on the part of the laboratory or by DGS-DCLS.

B. If a laboratory’s response is unacceptable and DGS-DCLS does not require additional on-site assessments, the laboratory shall complete another study. Such additional studies shall be conducted at least 15 calendar days from the previous study until the results are acceptable to DGS-DCLS. DGS-DCLS may conduct additional on-site assessments as necessary based on the results of any additional studies.

C. When the DMR-QA whole effluent toxicity portion does not include all test procedures required for a permit, the laboratory shall perform a proficiency test for aquatic toxicity testing.

D. DGS-DCLS shall not base loss of certification for aquatic toxicity testing solely on PT results. ]

1VAC30-45-540. (Reserved.)

1VAC30-45-550. (Reserved.)

1VAC30-45-560. (Reserved.)

1VAC30-45-570. (Reserved.)

1VAC30-45-580. (Reserved.)

1VAC30-45-590. (Reserved.)

Article 4 
Quality System

1VAC30-45-600. Quality system.

A. This article sets out the general requirements that an environmental laboratory has to successfully demonstrate to be recognized as competent to carry out specific environmental tests. The environmental laboratory shall establish, implement and maintain a quality system based on the required elements contained in this article.

B. The quality system shall be appropriate to the type, range and volume of testing, analysis, measurement or monitoring performed by the laboratory. [ Therefore, for technical or other reasons, some of the requirements of this article may not apply to every laboratory subject to this chapter. When in doubt as to the applicability of an Article 4 requirement, the applicant laboratory should consult DGS-DCLS. ]

C. If more stringent standards or requirements are included in a mandated test method or by regulation, the laboratory shall demonstrate that such requirements are met. If it is not clear which [ requirements are standard or requirement is ] more stringent, the standard [ or requirement ] from the method or regulation is to be followed.

D. Provisions pertaining to the management of the quality system appear in 1VAC30-45-610 through 1VAC30-45-700. Provisions pertaining to the technical requirements for the quality system appear in 1VAC30-45-710 through 1VAC30-45-770.

1VAC30-45-610. Quality manual.

A. General.

1. The laboratory shall document its quality system in a quality manual. The quality manual shall reflect all quality assurance and quality control practices and programs used by the laboratory. The required elements of the quality system may be described in more than one document.

2. The quality manual shall be maintained current under the responsibility of the quality assurance officer.

3. The quality manual and any related documents shall be communicated to, understood by, available to, and implemented by all laboratory personnel.

4. The quality manual shall include but not be limited to the elements listed in subsection B of this section.

B. The elements of a quality manual shall include but not be limited to:

1. Title page. The quality manual shall list the following items on the title page:

a. A document title;

b. The laboratory’s full name and address;

c. The name, address (if different from above), and telephone number of the responsible official, laboratory manager, and quality assurance officer;

d. The laboratory facility or facilities covered by the quality manual;

e. Signed and dated concurrence, with appropriate titles, of the responsible official, laboratory manager, and quality assurance officer; and

f. The effective date of the quality manual.

2. Table of contents.

3. A quality policy statement, including objectives [ of the quality system ] and [ commitments, signed by top management commitment to good laboratory practices ].

4. The organization and management structure of the laboratory, its place in any parent organization and relevant organizational charts.

5. The relationship between management, technical operations, support services and the quality system.

6. The capabilities of the laboratory or scope of its operation.

[ 7. Job descriptions of key staff and reference to the job descriptions of other staff.

8. Processes or procedures for establishing that personnel have adequate training and experience in the duties they are expected to carry out and are receiving any needed training.

9. Ethics policy statement developed by the laboratory. Processes and procedures for educating and training personnel in their ethical and legal responsibilities including the potential penalties for improper, unethical or illegal actions.

10. Mechanisms for ensuring that the laboratory reviews all new work to ensure that it has the appropriate facilities and resources before commencing such work. ]

[ 7. 11. ] Procedures to ensure that all records required by this chapter are retained, as well as procedures for control and maintenance of documentation through a document control system that ensures that all standard operating procedures, manuals, or documents clearly indicate the time period during which the procedure or document was in force.

[ 8. Job descriptions of key staff and reference to the job descriptions of other staff.

12. Procedures for dealing with complaints.

13. Procedures for audits and data review.

14. Reference to verification practices that may include interlaboratory comparisons, proficiency testing programs, use of reference materials and internal quality control schemes.

15. Procedures to be followed for feedback and corrective action whenever testing discrepancies are detected, or departures from documented policies and procedures occur.

16. The laboratory management arrangements for permitting departures from documented policies and procedures or from standard specifications when the departures are planned and controlled.

17. Reference to the major equipment and reference measurement standards used as well as the physical facility and environment used by the laboratory in conducting tests.

18. Reference to procedures for calibration, verification and maintenance of equipment.

19. A list of all technology/methods under which the laboratory performs its certified testing. ]

[ 9. 20. ] The laboratory's procedures for achieving traceability of measurements, including standards.

[ 10. A list of all test methods under which the laboratory performs its certified testing.

11. Mechanisms for ensuring that the laboratory reviews all new work to ensure that it has the appropriate facilities and resources before commencing such work.

12. Reference to the calibration and verification test procedures used.

13. 21. ] Procedures for receiving, handling, storing, and disposing of submitted samples.

[ 14. Reference to the major equipment and reference measurement standards used as well as the physical facility and environment used by the laboratory in conducting tests.

15. Reference to procedures for calibration, verification and maintenance of equipment.

16. Reference to verification practices, which may include interlaboratory comparisons, proficiency testing programs, use of reference materials and internal quality control schemes.

17. Procedures to be followed for feedback and corrective action whenever testing discrepancies are detected, or departures from documented policies and procedures occur.

18. The laboratory management arrangements for permitting departures from documented policies and procedures or from standard specifications when the departures are planned and controlled.

19. Procedures for dealing with complaints.

20. Procedures for audits and data review.

21. Processes or procedures for establishing that personnel have adequate training and experience in the duties they are expected to carry out and are receiving any needed training.

22. Ethics policy statement developed by the laboratory. Processes and procedures for educating and training personnel in their ethical and legal responsibilities including the potential penalties for improper, unethical or illegal actions.

23. 22. ] Reference to procedures for reporting analytical results.

C. Review and approval of quality manual.

1. The quality assurance officer shall review the laboratory’s quality assurance program, manual and any related documentation whenever there is any change in test methods employed by the laboratory, change in equipment, or any other change in the laboratory that [ may significantly affect affects ] the quality assurance program.

2. The quality assurance manual shall be reviewed and approved by the quality assurance officer, the laboratory manager, and the responsible official at least annually.

1VAC30-45-620. Organization.

The laboratory shall specify and document the functional responsibility, level of authority, and interrelationship or lines of communication of all personnel who manage, perform or verify work affecting the quality of tests, analyses, measurements and monitoring. One person may cover more than one organizational function. Each manager and employee of the laboratory shall have a clear understanding of his or her duties and responsibilities and the relationship of those responsibilities to the overall work of the laboratory.

1VAC30-45-630. Records.

The laboratory shall maintain a record system to suit its particular circumstances and comply with any applicable regulations. [ The This ] system shall produce unequivocal, accurate records [ which that ] document all laboratory activities.

1VAC30-45-640. Recordkeeping system and design.

A. The [ laboratory shall have a ] recordkeeping system [ shall allow that allows ] historical reconstruction of all laboratory activities that produced the analytical data.  The history of the sample shall be readily understood through the documentation. This shall include interlaboratory transfers of samples or extracts or both.

B. The records shall include the identity of personnel involved in sampling, [ sample preservation, ] sample receipt, preparation, calibration or testing [ , all documentation sent by the person transmitting the sample, including a chain of custody record form, if utilized ].

C. [ All The laboratory shall document all ] information relating to the laboratory [ facilities facility's ] equipment, analytical test methods, and related laboratory activities, such as sample receipt, sample preparation, or data verification [ shall be documented ].

D. The [ laboratory shall have a ] recordkeeping system [ shall facilitate that facilitates ] the retrieval of all working files and archived records for inspection and verification purposes, e.g., set format for naming electronic files.

E. [ All Responsible staff shall sign or initial all ] changes to records [ shall be signed or initialed by responsible staff ]. The reason for the signature or initials shall be clearly indicated in the records such as "sampled by," "prepared by," or "reviewed by."

F. All generated data except those that are generated by automated data collection systems, shall be recorded directly, promptly and legibly in permanent ink.

G. Entries in records shall not be obliterated by methods such as erasures, overwritten files or markings. All corrections to recordkeeping errors shall be made by one line marked through the error.  The individual making the correction shall sign or initial and date the correction. These criteria also shall apply to electronically maintained records.

H. [ Computer The laboratory shall keep computer ] and electronic data records [ shall be kept ] in accordance with [ the pertinent provisions of this section, ] 1VAC30-45-650 C [ and E ] and 1VAC30-45-730 K.

1VAC30-45-650. Records management and storage.

A. [ All The laboratory shall keep all ] records, certificates and reports [ shall be kept ] as required by applicable state and federal recordkeeping laws and regulations [ and. The laboratory shall  ] safely [ stored store these records ] and [ held hold them ] secure.

B. [ All The laboratory shall retain all ] records [ shall be retained ] for a minimum of three years from generation of the last entry in the records, or longer, if required by an applicable regulatory program, whichever is greater. [ All The laboratory shall maintain all ] information necessary for the historical reconstruction of data, including all original observations, calculations and derived data, calibration records and a copy of the test report [ , shall be maintained by the laboratory ].

C. Records [ which that ] are stored only on electronic media shall be supported by the hardware and software necessary for their retrieval.  Records that are stored or generated by computers or personal computers shall have hard copy or write-protected backup copies.

D. The laboratory shall establish a record management system for control of laboratory notebooks, instrument logbooks, standards logbooks, and records for data reduction, validation storage and reporting.

E. Access to archived information shall be documented with an access log. [ These The laboratory shall protect these ] records [ shall be protected ] against fire, theft, loss, environmental deterioration, vermin and, in the case of electronic records, electronic or magnetic sources.

F. The laboratory shall have a plan to ensure that the records are maintained or transferred in the event that a laboratory transfers ownership or goes out of business.  In addition, in cases of bankruptcy, [ the laboratory shall follow ] appropriate regulatory and state legal requirements concerning laboratory records [ shall be followed ] .

1VAC30-45-660. Required records.

A. Sample handling.

1. [ A The laboratory shall maintain a ] record of all procedures to which a sample is subjected while in the possession of the laboratory [ shall be maintained ]. These shall include but are not limited to all records pertaining to sample preservation, identification, receipt, acceptance or rejection, log-in, storage and tracking. [ The laboratory shall also maintain sampling information on each sample. This includes time and date of collection, type of sample (grab or composite), type of container, sampling point and preservation. ]

2. The laboratory shall have documented procedures for the receipt and retention of [ test items samples, including provisions necessary to protect the integrity of the samples ] .

B. Laboratory support activities. The [ laboratory shall retain the ] following documents and data [ shall be retained ]:

1. All original raw data, whether hard copy or electronic, for calibrations, samples and quality control measures, including analysts’ work sheets and data output records (chromatograms, strip charts, and other instrument response readout records).

2. A written description or reference to the specific test method used that includes a description of the specific computational steps used to translate parametric observations into a reportable analytical value.

3. Copies of final reports.

4. Archived standard operating procedures.

5. Correspondence relating to laboratory activities.

6. All corrective action reports, audits and audit responses.

7. Proficiency test results and raw data.

8. Results of data review, verification, and cross-checking procedures.

C. Analytical records. [ Essential The laboratory shall retain essential ] information associated with analytical documents, such as strip charts, tabular printouts, computer data files, analytical notebooks, and run logs [ , shall be retained ].  This information includes, but is not limited to, all manual calculations, e.g., manual integrations; sample preparation; standard and reagent origin, receipt, preparation, and use; quality control protocols and assessment; and method performance criteria.

D. Administrative records.  The [ laboratory shall maintain the ] following administrative records [ shall be maintained ]:

1. Personnel qualifications, experience and training records.

2. Records of demonstration of capability  for each analyst [ or work cell ].

3. A log of names, initials and signatures for all individuals who are responsible for signing or initialing any laboratory record.

1VAC30-45-670. Audits.

A. Internal audits.

1. The laboratory shall arrange for annual internal audits to verify that its operations continue to comply with the requirements of the laboratory’s quality system.  It is the responsibility of the quality assurance officer to plan and organize audits as required by a predetermined schedule and requested by management.

2. [ Such audits shall be carried out by trained and qualified personnel Trained ] and qualified personnel who are, wherever resources permit, independent of the activity to be audited [ , shall carry out these audits ]. Personnel shall not audit their own activities except when it can be demonstrated that an effective audit will be carried out.

3. Where the audit findings cast doubt on the correctness or validity of the laboratory's calibrations or test results, the laboratory shall take immediate corrective action.

4. [ Small laboratories A laboratory ] may have an audit performed under contract by an outside source competent to audit the laboratory’s operations.

B. Managerial review.

1. The laboratory management shall conduct a review, at least annually, of its quality system and its testing and calibration activities to ensure its continuing suitability and effectiveness and to introduce any necessary changes or improvements in the quality system and laboratory operations.

2. The review shall take account of reports from managerial and supervisory personnel, the outcome of recent internal audits, assessments by external bodies, the results of interlaboratory comparisons or proficiency tests, corrective actions and other relevant factors.

3. The laboratory shall have a procedure for review by management and maintain records of review findings and actions.

4. Where the staff of a laboratory is limited to a single analyst, a supervisor may perform a managerial review.

C. Audit review.  All audit and review findings and any corrective actions that arise from them shall be documented.  The laboratory management shall ensure that these actions are discharged within the agreed time frame as indicated in the quality manual or standard operating procedures or both.  For clarification, documentation of audit and review findings should be a simple procedure, essentially a memorandum setting out the findings of the audit and managerial review and any action to follow.

[ D. Performance audits.  In addition to periodic audits, the laboratory shall ensure the quality of results by implementing checks to monitor the quality of the laboratory’s analytical activities.  The following are examples of such checks:

1. Internal quality control procedures using statistical techniques.

2. Participation in proficiency testing or other interlaboratory comparisons.

3. Use of certified reference materials or in-house quality control using secondary reference materials.

4. Replicate testings using the same or different test methods.

5. Retesting of retained samples.

6. Correlation of results for different but related analysis of a sample (for example, total phosphorus should be greater than or equal to orthophosphate).

E. D. ] Corrective actions.

1. In addition to providing acceptance criteria and specific protocols for corrective actions in the method standard operating procedures, the laboratory shall implement general procedures to be followed to determine consistently when departures from documented policies, procedures and quality control have occurred.  These procedures may include but are not limited to the following:

a. Identify the individual or individuals responsible for assessing each  quality control data type;

b. Identify the individual or individuals responsible for initiating or recommending corrective actions or both;

c. Define how the analyst shall treat a data set if the associated quality control measurements are unacceptable;

d. Specify how out-of-control situations and subsequent corrective actions are to be documented; and

e. Specify procedures for management (including the quality assurance officer) to review corrective action reports.

2. To the extent possible, samples shall be reported only if all quality control measures are acceptable.  If a quality control measure is found to be out of control, and the data are to be reported, all samples associated with the failed quality control measure shall be reported with the appropriate data qualifiers.

1VAC30-45-680. Subcontracting analytical samples.

A. Where a laboratory subcontracts any part of the testing covered under this chapter, the testing shall only be subcontracted to a laboratory certified under 1VAC30-46 [ or under another state’s NELAP-approved program ].

B. The report from the subcontractor shall be a separate part of the laboratory report and identified as laboratory testing done by a subcontractor.

C. The laboratory shall retain records demonstrating that the requirements of this section have been met.

1VAC30-45-690. Outside support services and supplies.

[ A. Where the laboratory procures outside services and supplies in support of tests, the laboratory shall use only those outside support services and supplies that are of adequate quality to sustain confidence in the laboratory's tests.

B. Where no independent assurance of the quality of outside support services or supplies is available, the A. The ] laboratory shall have [ a policy and ] procedures [ to ensure that purchased equipment, materials and services comply with specified requirements for the selection and purchasing of services and supplies it uses that affect the quality of environmental tests. Procedures shall exist for the purchase, reception and storage of reagents and laboratory consumable materials relevant for the environmental tests ].

[ B. ] The laboratory shall ensure that purchased equipment and consumable materials are not used until they have been inspected, calibrated or otherwise verified as complying with [ any ] standard specifications [ relevant to the calibrations or or requirements defined in the methods for the ] tests concerned. [ These services and supplies used shall comply with specified requirements. Records of actions taken to check compliance shall be maintained. ]

C. The laboratory shall maintain records of all suppliers from whom it obtains support services or supplies required for tests.

1VAC30-45-700. Complaints.

The laboratory shall have documented policy and procedures for the resolution of complaints about the laboratory's activities. Where a complaint or any other circumstance raises doubt concerning the laboratory's compliance with the laboratory's policies or procedures, or with the requirements of this chapter or otherwise concerning the quality of the laboratory's calibrations or tests, the laboratory shall ensure that those areas of activity and responsibility involved are promptly audited in accordance with 1VAC30-45-670 A. Records of the complaint and subsequent actions shall be maintained.

1VAC30-45-710. Environment and work areas.

Laboratory accommodations, test areas, energy sources, lighting, heating and ventilation shall be such as to facilitate proper performance of tests.  Laboratories [ may shall ]  meet the requirements of subdivisions 1 through 8 of this section as appropriate to provide compliance with this requirement.

1. The environment in which these activities are undertaken shall not invalidate the results or adversely affect the required accuracy of measurement.  Particular care shall be taken when such activities are undertaken at sites other than the permanent laboratory premises.

2. The laboratory shall provide for the effective monitoring, control and recording of environmental conditions as appropriate.  Such environmental conditions may include biological sterility, dust, electromagnetic interference, humidity, mains voltage, temperature, and sound and vibration levels.

3. In instances where monitoring or control of any of the above-mentioned items are specified in a test method or by regulation, the laboratory shall meet and document adherence to the laboratory facility requirements.

4. There shall be effective separation between [ testing neighboring ] areas [ when the activities in the testing areas in which there ] are incompatible [ (i.e., microbiological culture or incubation and volatile organic chemicals) activities including culture handling or incubation areas and volatile organic chemical handling areas. The laboratory shall take measures to prevent cross-contamination ] .

5. Access to and use of all areas affecting the quality of these activities shall be defined and controlled.

6. Adequate measures shall be taken to ensure good housekeeping in the laboratory and to ensure that any contamination does not adversely affect data quality.

7. Work spaces shall be available to ensure an unencumbered work area.

8. Work areas include:

a. Access and entryways to the laboratory;

b. Sample receipt areas;

c. Sample storage areas;

d. Chemical and waste storage areas; and

e. Data handling and storage areas.

1VAC30-45-720. Equipment and reference materials.

A. The laboratory shall be furnished with all items of equipment, including reference materials, required for the correct performance of tests for which certification is sought.  The laboratory shall maintain records of reference materials sufficient to provide proper performance of tests.  In those cases where the laboratory needs to use equipment outside its permanent control it shall ensure that the relevant requirements of this article are met.

B. All equipment shall be properly maintained, inspected and cleaned.  Maintenance procedures shall be documented.

C. Any item of the equipment that has been subjected to overloading or mishandling, [ which or that ] gives suspect results, or has been shown by verification or otherwise to be defective shall be taken out of service [ immediately ], clearly identified [ as being out of service ] and, wherever possible, stored at a specified place until it has been repaired and shown by calibration, verification or test to perform satisfactorily. The laboratory shall examine the effect of this defect on previous calibrations or tests.

D. Each item of equipment including reference materials shall be labeled, marked or otherwise identified to indicate its calibration status.

E. Records of each major item of equipment significant to the tests performed shall be maintained. These records shall include documentation on all routine and nonroutine maintenance activities. The laboratory shall maintain records of reference materials sufficient to provide proper performance of tests. The records [ may shall ] include:

1. The name of the item of equipment;

2. The manufacturer's name, type identification, and serial number or other unique identification;

3. Date received and date placed in service (if available);

4. Current location, where appropriate;

5. If available, condition when received (e.g., new, used, reconditioned);

6. Copy of the manufacturer's instructions, where available;

7. Dates and results of calibrations or verifications or both and date of the next calibration or verification;

8. Details of maintenance carried out to date and planned for the future; and

9. History of any damage, malfunction, modification or repair.

1VAC30-45-730. Test methods and standard operating procedures.

A. Methods documentation.

1. The laboratory shall have documented instructions on the use and operation of all relevant equipment, on the handling and preparation of samples, and for calibration or testing, where the absence of such instructions could jeopardize the calibrations or tests.

2. All instructions, standards, manuals and reference data relevant to the work of the laboratory shall be maintained up to date and be readily available to the staff.

B. Standard operating procedures (SOPs).

1. Laboratories shall maintain SOPs that accurately reflect all phases of current laboratory activities such as assessing data integrity, corrective actions, handling customer complaints, and all test methods. These documents, for example, may be equipment manuals provided by the manufacturer or internally written documents. The test methods may be copies of published methods as long as any changes or selected options in the methods are documented and included in the laboratory methods manual.

2. The SOPs shall be organized. Each SOP shall clearly indicate the effective date of the document, the revision number, and the signature or signatures of the [ approving authority responsible laboratory manager or managers ].

3. Copies of all SOPs shall be accessible to all personnel.

C. Laboratory methods manuals.

1. The laboratory shall have and maintain an in-house methods manual or manuals for each certified analyte or test method.

2. This manual may consists of copies of published or referenced methods or standard operating procedures that have been written by the laboratory. In cases where modifications to the published method have been made by the laboratory or where the referenced test method is ambiguous or provides insufficient detail, these changes or clarifications shall be clearly described. Each test method shall include or reference where applicable:

a. Identification of the test method;

b. Applicable matrix or matrices;

c. Method detection limit;

d. Scope and application, including components to be analyzed;

e. Summary of the test method;

f. Definitions;

g. Interferences;

h. Safety;

i. Equipment and supplies;

j. Reagents and standards;

k. Sample collection, preservation, shipment and storage;

l. Quality control;

m. Calibration and standardization;

n. Procedure;

o. Calculations;

p. Method performance;

q. Pollution prevention;

r. Data assessment and acceptance criteria for quality control measures;

s. Corrective actions for out-of-control data;

t. Contingencies for handling out-of-control or unacceptable data;

u. Waste management;

v. References; and

w. Any tables, diagrams, flowcharts and validation data.

D. Test methods.

1. Laboratories shall use (i) promulgated test methods in accordance with the Code of Federal Regulations; (ii) test methods stated in any current permit issued by Virginia Air Pollution Control Board, the Virginia Waste Management Board, or the State Water Control Board; or (iii) alternate test procedures approved by the board issuing the permit or the Department of Environmental Quality, including applicable quality assurance requirements, and sample preservation, container, storage, and holding time requirements.

2. The laboratory shall use appropriate test methods and procedures for all tests and related activities within its responsibility (including sample handling, transport and storage, preparation and analysis).  The method and procedures shall be consistent with the accuracy required and with any standard specifications relevant to the calibrations or tests concerned.

3. When the use of reference test methods for a sample analysis is mandated, only those methods shall be used.

4. Where test methods are employed that are not required, as in the Performance Based Measurement System approach, the methods shall be fully documented and validated (see subsection E of this section).

E. Demonstration of capability.

1. Prior to acceptance and institution of any test method, satisfactory demonstration of method capability is required. In general, this demonstration does not test the performance of the method in real world samples, but in the applicable and available clean [ quality system ] matrix [ sample ] (a [ sample of a quality system ] matrix in which no target analytes or interferences are present at concentrations that impact the results of a specific test method), e.g., [ drinking ] water, solids, biological tissue and air. [  For analytes that do not lend themselves to spiking, the demonstration of capability may be performed using quality control samples Laboratories shall follow the procedure in subsection F of this section to demonstrate capability ].

2. Thereafter, continuing demonstration of method performance, such as laboratory control samples, is required.

3. In cases where a laboratory analyzes samples using a test method that has been in use by the laboratory before July 1999, and there have been no significant changes in instrument type, personnel or test method, the continuing demonstration of method performance and the analyst’s documentation of continued proficiency shall be acceptable. The laboratory shall have records on file to demonstrate that an initial demonstration of capability is not required.

4. In all cases, [ the laboratory shall complete and retain ] a certification statement [ shall be completed and retained by the laboratory to be made and shall make the statement ] available upon request. [ All The laboratory shall retain all ] associated supporting data necessary to reproduce the analytical results summarized in the certification statement [ shall be retained by the laboratory ].

5. [ A The laboratory shall complete a ] demonstration of capability [ shall be completed ] each time there is a change in instrument type, personnel or test method.

[ 6. In laboratories with specialized work cells (a group consisting of analysts with specifically defined tasks that together perform the test method), the group as a unit shall meet the criteria of this subsection. This demonstration of capability shall be fully documented. ]

F. Procedure for demonstration of capability. [ 1. ]  The following steps [ (adapted from the EPA test methods published in 40 CFR Part 136, Appendix A) ] shall be performed [ if required reagents/standards are available for mandated test methods ].  [ However, before any results are reported using this method, actual sample spike results may be used to meet this standard, i.e., at least four consecutive matrix spikes within the last 12 months. For analytes that do not lend themselves to spiking, e.g. TSS, the demonstration of capability may be performed using quality control samples. The laboratory may document that other approaches to demonstration of capability are adequate. This documentation shall be included in the laboratory’s quality manual. ]:

[ a. 1. ] A quality control (QC) sample [ shall may ] be obtained from an outside source [ .  If not available, the QC sample or ] may be prepared by the laboratory using [ alternate source ] stock standards that are prepared independently from those used in instrument calibration.

[ b. 2. ] The analyte or analytes shall be diluted in a volume of clean [ quality system ] matrix sufficient to prepare four aliquots at the concentration specified [ , ] or [ , ] if unspecified, to a concentration [ approximately 10 times the method-stated or laboratory-calculated method detection limit of 1-4 times the limit of quantitation ].

[ c. 3. ] At least four aliquots shall be prepared and analyzed according to the test method either concurrently or over a period of days.

[ d. 4. ] Using all of the results, calculate the mean recovery in the appropriate reporting units (such as g/L) and the standard deviations of the population sample (n-1)  (in the same units) for each parameter of interest.  When it is not possible to determine mean and standard deviations, such as for presence or absence of the analyte and logarithmic values, the laboratory [ must  shall ] assess performance against established and documented criteria.

[ e. 5. ]  Compare the information from subdivision [ 1 d 4 ] of this subsection to the corresponding acceptance criteria for precision and accuracy in the test method (if applicable) or in laboratory-generated acceptance criteria (if there are [ no not ] established mandatory criteria).  If all parameters meet the acceptance criteria, the analysis of actual samples may begin.  If any one of the parameters do not meet the acceptance criteria, the performance is unacceptable for that parameter.

[ f. 6. ] When one or more of the tested parameters fail at least one of the acceptance criteria, the analyst [ must shall ] proceed according to either subdivision [ (1) (a) or [ (2) (b) ] below.

[ (1) (a) ] Locate and correct the source of the problem and repeat the test for all parameters of interest beginning with subdivision [ 1 c 3 ] of this subsection.

[ (2) (b) ] Beginning with subdivision [ 1 c 3 ] of this subsection, repeat the test for all parameters that failed to meet criteria.  Repeated failure, however, confirms a general problem with the measurement system.  If this occurs, locate and correct the source of the problem and repeat the test for all compounds of interest beginning with subdivision [ 1 c 3 ] of this subsection.

[ 2. It is the responsibility of the laboratory to document that other approaches to demonstrating capability are adequate.  This documentation shall be included in the laboratory’s quality assurance manual. ]

G. Certification statement.  The following certification statement shall be used to document the completion of each demonstration of capability.  A copy of the certification statement shall be retained in the personnel records of each affected employee.

Demonstration of Capability

Certification Statement

Date: Page __of __

Laboratory Name:

Laboratory Address:

Analyst(s) Name(s):

Matrix:

(examples: laboratory pure water, soil, air, solid, biological  tissue)

Method number, SOP#, Rev#, and Analyte, or Class of Analytes or Measured Parameters

(examples: barium by 200.7, trace metals by 6010 [ B ], benzene by 8021 [ B ], etc.)

We, the undersigned, CERTIFY that:

1. The analysts identified above, using the cited test method(s), which is in use at this facility for the analyses of samples under the Virginia Environmental Laboratory Certification Program, have met the Demonstration of Capability.

2. The test method(s) was performed by the analyst(s) identified on this certification.

3. A copy of the test method(s) and the laboratory-specific SOPs are available for all personnel on-site.

4. The data associated with the demonstration capability are true, accurate, complete and self-explanatory (1).

5.  All raw data (including a copy of this certification form) necessary to reconstruct and validate these analyses have been retained at the facility, and that the associated information is well organized and available for review by authorized assessors.

_________________________________________________

Laboratory Manager’s Name and Title       Signature       Date

_________________________________________________

Quality Assurance Officer’s Name              Signature       Date

[ (1) True - consistent with supporting data. Accurate - based on good laboratory practices consistent with sound scientific principles and practices. Complete - includes the results of all supporting performance testing. Self-explanatory - data properly labeled and stored so that the results are clear and require no additional explanation. ]

H. Sample aliquots. Where sampling (as in obtaining sample aliquots from a submitted sample) is carried out as part of the test method, the laboratory shall use documented procedures and appropriate techniques to obtain representative subsamples.

I. Data verification. Calculations and data transfers shall be subject to appropriate checks. The laboratory shall establish standard operating procedures to ensure that (i) the reported data are free from transcription and calculation errors and (ii) all quality control measures are reviewed and evaluated before data are reported.  The laboratory also shall establish standard operating procedures addressing manual calculations including manual integrations.

J. Documentation and labeling of standards and reagents. Documented procedures shall exist for the reception and storage of consumable materials used for the technical operations of the laboratory.

1. The laboratory shall retain records for all standards, reagents [ , reference materials ] and media including the manufacturer/vendor, the manufacturer’s Certificate of Analysis or purity (if [ supplied available ]), the date of receipt, recommended storage conditions, and an expiration date after which the material shall not be used unless its reliability is verified by the laboratory.

2. Original containers (such as provided by the manufacturer or vendor) shall be labeled with an expiration date.

3. Records shall be maintained on [ reagent and ] standard [ and reference material ] preparation. These records shall indicate traceability to purchased stocks or neat compounds, reference to the method of preparation, date of preparation, expiration date and preparer's initials.

4. Sufficient identification of containers of prepared reagents and standards shall be provided to ensure proper performance of tests.

K. Computers and electronic data related requirements.  Where computers, automated equipment or microprocessors are used for the capture, processing, manipulation, recording, reporting, storage or retrieval of test data, the laboratory shall ensure the following:

[ 1. All requirements of this article are complied with.

2. 1. ] Computer software [ developed by the user ] is [ tested and ] documented [ to be in sufficient detail and is suitably validated as being ] adequate for use [ , e.g., internal audits, personnel training, focus point of quality assurance and quality control ].

[ 3. 2. ] Procedures are established and implemented for protecting the integrity of data, such as integrity of data entry or capture, data storage, data transmission and data processing.

[ 4. 3. ] Computer and automated equipment are maintained to ensure proper functioning and provided with the environmental and operating conditions necessary to maintain the integrity of calibration and test data.

[ 5. 4. ] Appropriate procedures are established and implemented for the maintenance of security of data including the prevention of unauthorized access to, and the unauthorized amendment of, computer records.

1VAC30-45-740. Measurement traceability and calibration.

A. General requirements. All [ measuring operations and testing ] equipment [ used for environmental tests, including equipment for subsidiary measurements (e.g., for environmental conditions) ] having [ an a significant ] effect on the accuracy or validity [ of tests of the result of the environmental test or sampling ] shall be calibrated [ or verified or both ] before being put into service and on a continuing basis. The laboratory shall have an established program [ and procedure ] for the calibration [ and verification ] of its [ measuring and test ] equipment. This includes balances, thermistors, thermometers and control standards. [ Such a program shall include a system for selecting, using, calibrating, checking, controlling and maintaining measurement standards, reference materials used as measurement standards, and measuring and test equipment used to perform environmental tests. ]

B. Traceability of calibration.

[ 1. The laboratory shall ensure that the equipment used can provide the uncertainty of measurement needed. ]

[ 1. 2. ]  The overall program of calibration or verification or both and validation of equipment shall be designed and operated so as to ensure that measurements made by the laboratory are traceable to national standards of measurement [ where available ].

[ 2. Calibration certificates shall indicate the traceability to national standards of measurement and shall provide the measurement results and associated uncertainty of measurement or a statement of compliance with an identified metrological specification or both.  The laboratory shall maintain records of all such certifications. ]

3. Where traceability [ of measurements ] to [ national standards of measurement the International System of Units (SI) ] is not [ applicable the possible or not relevant, the same requirements for traceability to, for example, certified reference materials, agreed methods and/or consensus standards, are required. ] [ , the The ] laboratory shall provide satisfactory evidence of correlation of results, for example by participation in a suitable program of interlaboratory comparisons, proficiency testing, or independent analysis.

C. Reference standards [ and reference materials ].

1. Reference [ standards. The laboratory shall have a program and procedure for the calibration of its reference standards. Reference standards of measurement shall be calibrated by a body that can provide traceability as described in subsection B of this section. Such reference  ] standards of measurement held by the laboratory (such as Class S or equivalent weights or traceable thermometers) shall be used for calibration only [ and for no other purpose ] , unless it can be demonstrated that their performance as reference standards [ have would ] not [ been be ] invalidated. [ Reference standards of measurement shall be calibrated by a body that can provide traceability.  Where possible, this traceability shall be to a national standard of measurement Where commercially available, this traceability shall be to a national standard of measurement ].

[ 2. There shall be a program of calibration and verification for reference standards.

3. Where relevant, reference standards and measuring and testing equipment shall be subjected to in-service checks between calibrations and verifications.  2. Reference materials. Reference materials shall [ , where commercially available, ] be traceable [ to SI units of measurement, or to certified reference materials ].  Where possible, traceability shall be to national or international standards of measurement, or to national or international standard reference materials. [ Internal reference materials shall be checked as far as is technically and economically practicable. ]

D. Calibration. Calibration requirements are divided into two parts: (i) requirements for analytical support equipment and (ii) requirements for instrument calibration. In addition, the requirements for instrument calibration are divided into initial instrument calibration and continuing instrument calibration verification.

1. Support equipment. These standards apply to all devices that may not be the actual test instrument, but are necessary to support laboratory operations.  These include but are not limited to balances, ovens, refrigerators, freezers, incubators, water baths, temperature measuring devices (including thermometers and thermistors), thermal/pressure sample preparation devices and volumetric dispensing devices (such as Eppendorf®, or automatic dilutor or dispensing devices) if quantitative results are dependent on their accuracy, as in standard preparation and dispensing or dilution into a specified volume.

a. All support equipment shall be maintained in proper working order.  The records of all repair and maintenance activities, including service calls, shall be kept.

b. All support equipment shall be calibrated or verified at least annually, using NIST traceable references when available, over the entire range of use.  The results of such calibration shall be within the specifications required of the application for which this equipment is used.  If not, the laboratory shall either (i) remove the equipment from service until repaired or (ii) maintain records of established correction factors to correct all measurements.

c. Raw data records shall be retained to document equipment performance.

d. Prior to use on each working day, balances, ovens, refrigerators, freezers, and water baths shall be checked in the expected use range, with NIST traceable references where available.  The acceptability for use or continued use shall be according to the needs of the analysis or application for which the equipment is being used.

e. Mechanical volumetric dispensing devices including burettes (except Class A glassware) shall be checked for accuracy on at least a quarterly use basis.  Glass microliter syringes are to be considered in the same manner as Class A glassware, but shall come with a certificate attesting to established accuracy or the accuracy shall be initially demonstrated and documented by the laboratory.

f. For chemical tests, the temperature, cycle time and pressure of each run of autoclaves shall be documented by the use of appropriate chemical indicators or temperature recorders and pressure gauges.

g. For biological tests that employ autoclave sterilization, the following requirements apply:

(1) The performance of each autoclave shall be initially evaluated by establishing its functional properties and performance, for example heat distribution characteristics with respect to typical uses.  Autoclaves shall meet specified temperature tolerances.  Pressure cookers fitted only with a pressure gauge are not recommended for sterilization of media or decontamination of wastes.

(2) Records of autoclave operations including temperature and time shall be maintained.  This shall be done for every cycle.  Acceptance and rejection criteria shall be established and used to evaluate the autoclave efficiency and effectiveness.

2. Instrument calibration.

a. This standard specifies the essential elements that define the procedures and documentation for initial instrument calibration and continuing instrument calibration verification to ensure that the data shall be of known quality and be appropriate for a given regulation or decision.  This standard does not specify detailed procedural steps for calibration, but establishes the essential elements for selection of the appropriate technique or techniques.  If more stringent standards or requirements are included in a mandated test method or by regulation, the laboratory shall demonstrate that such requirements are met.  If it is not apparent which standard is more stringent, then the requirements of the regulation or mandated test method are to be followed.

b. Initial instrument calibrations. The following items are essential elements of initial instrument calibration:

(1) The [ laboratory shall include or reference the ] details of the initial instrument calibration procedures, including calculations, integrations, acceptance criteria and associated statistics [ shall be included or referenced ] in the [ test method ] standard operating procedure [ for the test method ]. When initial instrument calibration procedures are referenced in the test method, then the [ laboratory shall retain the ] referenced material [ shall be retained by the laboratory ] and [ be make it ] available for review.

(2) [ Sufficient The laboratory shall retain sufficient ] raw data records [ shall be retained ] to permit reconstruction of the initial instrument calibration, e.g., calibration date, test method, instrument, analysis date, each analyte name, analyst’s initials or signature, concentration and response, calibration curve or response factor, or unique equation or coefficient used to reduce instrument responses to concentration.

(3) Sample results shall be quantitated from the initial instrument calibration and may not be quantitated from any continuing instrument calibration verification [ unless otherwise required by regulation, method, or program ] .

(4) All initial instrument calibrations shall be verified with a standard obtained from a second manufacturer or lot. Traceability shall be to a national standard, when available. This element does not apply to laboratories performing only simple test procedures.

(5) Criteria for the acceptance of an initial instrument calibration shall be established, e.g., correlation coefficient and relative percent difference. The criteria used shall be [ appropriate to the technique employed 0.995 or greater for the calibration coefficient unless a different criterion is included in the method being used ].

(6) Results of samples not bracketed by initial calibration standards (within calibration range) shall be reported as having less certainty, e.g., defined qualifiers or flags or explained in the case narrative. The lowest calibration standard shall be above the detection limit.

(7) If the initial instrument calibration results are outside established acceptance criteria, corrective actions shall be performed. Data associated with an unacceptable initial instrument calibration shall not be reported.

(8) Calibration standards shall include concentrations at or below the regulatory limit or decision level, if these limits or levels are known by the laboratory, unless these concentrations are below the laboratory’s demonstrated detection limits.

(9) If a reference or mandated method does not specify the number of calibration standards, the minimum number is two, not including blanks or a zero standard. The laboratory shall have a standard operating procedure for determining the number of points for establishing the initial instrument calibration.

c. Continuing instrument calibration verification.

(1) When an initial instrument calibration is not performed on the day of analysis, the validity of the initial calibration shall be verified prior to sample analyses by a continuing instrument calibration check with each analytical batch. This provision does not apply to laboratories performing only simple test procedures.

(2) The following items are essential elements of continuing instrument calibration verification:

(a) The [ laboratory shall include or reference the ] details of the continuing instrument calibration procedure, calculations and associated statistics [ shall be included or referenced ] in the [ test method ] standard operating procedure [ for the test method ].

[ (b) The laboratory shall verify calibration for each compound, element, or other discrete chemical species, except for multicomponent analytes such as Aroclors, Total Petroleum Hydrocarbons, or Toxaphene where a representative chemical related substance or mixture can be used. ]

[ (b) A (c) The laboratory shall perform a ] continuing instrument calibration [ check shall be repeated at verification as follows:

(1) At ] the beginning and end of each analytical batch. [ The concentrations of the calibration verification shall be varied within the established calibration range. ] If an internal standard is used, only one [ continuing instrument calibration ] verification [ shall be analyzed per analytical batch needs to be performed at the beginning of the ] analytical batch;

[ (2) Whenever it is expected that the analytical system may be out of calibration or might not meet the verification acceptance criteria;

(3) If the time period for calibration or the most previous calibration verification has expired; or

(4) For analytical systems that contain a calibration verification requirement. ]

[ (c) (d) ] Sufficient raw data records shall be retained to permit reconstruction of the continuing instrument calibration verification, e.g., or test method, instrument, analysis date, each analyte name, concentration and response, calibration curve or response factor, or unique equations or coefficients used to convert instrument responses into concentrations. Continuing calibration verification records shall explicitly connect the continuing verification data to the initial instrument calibration.

[ (d) (e) ] Criteria for the acceptance of a continuing instrument calibration verification shall be established, e.g., [ percent recovery or ] relative percent difference.

[ (e) (f) ] If the continuing instrument calibration verification results obtained are outside established acceptance criteria, corrective actions shall be performed.  If routine corrective action procedures fail to produce a second consecutive (immediate) calibration verification within acceptance criteria, then either the laboratory has to demonstrate [ acceptable ] performance after corrective action with two consecutive successful calibration verifications, or a new initial instrument calibration shall be performed.  If the laboratory has not [ demonstrated acceptable performance verified calibration ], sample analyses shall not occur until [ a new initial calibration curve is established and verified the analytical system is calibrated or calibration verified ].  [ However, sample data If samples are analyzed using a system on which the calibration has not yet been verified, the results shall be flagged. Data ] associated with an unacceptable calibration verification may be [ reported as qualified data fully useable ] under the following special conditions:

(i) When the acceptance criteria for the continuing calibration verification  are exceeded high, i.e., high bias, and there are associated samples that are nondetects, then those nondetects may be reported.  Otherwise the samples affected by the unacceptable calibration verification shall be reanalyzed after a new calibration curve has been established, evaluated and accepted.

(ii) When the acceptance criteria for the continuing calibration verification are exceeded low, i.e., low bias, those sample results may be reported if they exceed a maximum regulatory limit or decision level.  Otherwise the samples affected by the unacceptable verification shall be reanalyzed after a new calibration curve has been established, evaluated and accepted.

1VAC30-45-750. [ Essential quality control procedures Quality assurance ].

[ A. General. The laboratory shall have quality control procedures for monitoring the validity of environmental tests undertaken. The resulting data shall be recorded in such a way that trends are detectable and, where practicable, statistical techniques shall be applied to the reviewing of the results. This monitoring shall be planned and reviewed and may include, but not be limited to, the following:

1. Regular use of certified reference materials and/or internal quality control using secondary reference materials.

2. Participation in interlaboratory comparison or proficiency testing program

3. Replicate tests using the same or different methods.

4. Retesting of retained samples.

5. Correlation of results for different characteristics of a sample (for example, total phosphate should be greater than or equal to orthophosphate). ]

[ A. B. Essential quality control procedures. ] The general quality control principles in subsections [ B C ] through [ E F ] of this section shall apply, where applicable, to all environmental laboratories. The manner in which they are implemented is dependent on the types of tests performed by the laboratory. [ 1VAC30-45-760 through 1VAC30-45-829 specify quality control requirements for specific test types. ] The standards for any given test type shall assure that the applicable principles are addressed.

[ B. C. ] All laboratories shall have detailed written protocols in place to monitor the following quality controls:

1. Positive and negative controls to monitor tests such as blanks, spikes, reference toxicants.

2. Tests to define the variability or repeatability of the laboratory results or both such as replicates.

3. Measures to assure the accuracy of the test method including calibration or continuing calibrations or both, use of certified reference materials, proficiency test samples, or other measures.

4. Measures to evaluate test method capability, such as method detection limits and quantitation limits or range of applicability such as linearity.

5. Selection of appropriate formulae to reduce raw data to final results such as regression analysis, comparison to internal and external standard calculations, and statistical analyses.

6. Selection and use of reagents and standards of appropriate quality.

7. Measures to assure the selectivity of the test for its intended purpose.

8. Measures to assure constant and consistent test conditions (both instrumental and environmental) where required by the test method such as temperature, humidity, light, or specific instrument conditions.

[ C. All quality control measures shall be assessed and evaluated on an on-going basis, and quality control acceptance criteria shall be used to determine the validity of the data.

D. The laboratory shall have procedures for the development of acceptance or rejection criteria where no method or regulatory criteria exist. (See 1VAC30-45-760 B.)

E. The laboratory shall ensure that the essential quality control standards and protocols listed in subsection B of this section and specified by mandated methods or regulations are incorporated into the laboratory’s method manual and followed. ]

1VAC30-45-760. [ Quality control requirements.

A. General.

1. The quality control protocols specified by the laboratory’s method manual shall be followed  (1VAC30-45-730 C). The laboratory shall ensure that either the (i) applicable essential standards outlined in this section through 1VAC30-45-829 or (ii) mandated methods or regulations, whichever are more stringent, are incorporated into their method manuals. When it is not apparent which is more stringent, the quality control in the mandated method or regulations is to be followed.

2. All quality control measures shall be assessed and evaluated on an ongoing basis and quality control acceptance criteria shall be used to determine the validity of the data. The laboratory shall have procedures for the development of acceptance/rejection criteria where no method or regulatory criteria exists.

B. Initial test method evaluation. For all test methods other than toxicity and microbiology, the requirements of subdivisions 1 and 2 of this subsection apply. For toxicity and microbiology testing, the initial test method evaluation requirements are contained in 1VAC30-45-780 through 1VAC30-45-788 and 1VAC30-45-790 through 1VAC30-45-798, respectively. For the evaluation of precision and bias (subdivision 3 of this subsection), the requirements of subdivision 3 a of this subsection apply to standard methods. The requirements of subdivision 3 b of this subsection apply to the methods referenced in that subdivision.

1. Limit of detection (LOD).

a. The laboratory shall determine the LOD for the method for each target analyte of concern in the quality system matrices. All sample processing steps of the analytical method shall be included in the determination of the LOD.

b. The validity of the LOD shall be confirmed by qualitative identification of the analyte(s) in a quality control sample in each quality system matrix containing the analyte at no more than two to three times the LOD for single analyte tests and one to four times the LOD for multiple analyte tests. This verification shall be performed on every instrument that is to be used for analysis of samples and reporting of data.

c. An LOD study is not required for any component for which spiking solutions or quality control samples are not available such as temperature, or, when test results are not to be reported to the LOD (versus the limit of quantitation or working range of instrument calibration), according to 1VAC30-45-771, 1VAC30-45-805, 1VAC30-45-814, and 1VAC30-45-826. Where an LOD study is not performed, the laboratory may not report a value below the limit of quantitation.

2. Limit of quantitation (LOQ).

a. The laboratory shall determine the LOQ for each analyte of concern according to a defined, documented procedure.

b. The LOQ study is not required for any component or property for which spiking solutions or quality control samples are not commercially available or otherwise inappropriate (e.g., pH).

c. The validity of the LOQ shall be confirmed by successful analysis of a QC sample containing the analytes of concern in each quality system matrix one to two times the claimed LOQ. A successful analysis is one where the recovery of each analyte is within the established test method acceptance criteria or client data quality objectives for accuracy. This single analysis is not required if the bias and precision of the measurement system is evaluated at the LOQ.

3. Evaluation of precision and bias.

a. Standard methods. The laboratory shall evaluate the precision and bias of a standard method for each analyte of concern for each quality system matrix according to either of the following:

(1) The single-concentration four-replicate recovery study procedures in 1VAC30-45-730 F, or

(2) An alternate procedure documented in the quality manual when the analyte cannot be spiked into the sample matrix and quality control samples are not commercially available.

b. Nonstandard methods.

(1) For laboratory-developed test methods or nonstandard test methods that were not in use by the laboratory before July 2003, the laboratory shall have a documented procedure to evaluate precision and bias. The laboratory shall also compare results of the precision and bias measurements with criteria given in the reference method or criteria established by the laboratory.

(2) Precision and bias measurements shall evaluate the method across the analytical calibration range of the method. The laboratory shall also evaluate precision and bias in the relevant quality system matrices and shall process the samples through the entire measurement system for each analyte of interest.

(3) The following are examples of a systematic approach to evaluate precision and bias:

(a) Example 1. Analyze QC samples in triplicate containing the analytes of concern at or near the limit of quantitation, at the upper-range of the calibration (upper 20%) and at a mid-range concentration. Process these samples on different days as three sets of samples through the entire measurement system for each analyte of interest. Each day one QC sample at each concentration is analyzed. A separate method blank shall be subjected to the analytical method along with the QC samples on each of the three days. (Note that the three samples at the LOQ concentration can demonstrate sensitivity as well.) For each analyte, calculate the mean recovery for each day, for each level over days, and for all nine samples. Calculate the relative standard deviation for each of the separate means obtained. Compare the standard deviations for the different days and the standard deviations for the different concentrations. If the different standard deviations are all statistically insignificant (e.g., F-test), then compare the overall mean and standard deviation with the established criteria from above.

(b) Example 2. A validation protocol such as the Tier I, Tier II, and Tier III requirements in U.S. EPA Office of Water’s Alternate Test Procedure (ATP) approval process.

4. Evaluation of selectivity. The laboratory shall evaluate selectivity by following the checks established within the method. These checks may include mass spectral tuning, second column confirmation, ICP inter-element interference checks, chromatography retention time windows, sample blanks, spectrochemical absorption or fluorescence profiles, co-precipitation evaluations, and electrode response factors.

1VAC30-45-770. Chemical testing: positive and negative controls.

A. Negative control – method performance.

1. Purpose. The method blank is used to assess the preparation batch for possible contamination during the preparation and processing steps. The method blank shall be processed along with and under the same conditions as the associated samples to include all steps of the analytical procedure. Procedures shall be in place to determine if a method blank is contaminated. Any affected samples associated with a contaminated method blank shall be reprocessed for analysis or the results reported with appropriate data qualifying codes.

2. Frequency. The method blank shall be analyzed at a minimum of one per preparation batch. In those instances for which no separate preparation method is used (example: volatiles in water) the batch shall be defined as environmental samples that are analyzed together with the same method and personnel, using the same lots of reagents, not to exceed the analysis of 20 environmental samples.

3. Composition. The method blank shall consist of a quality system matrix that is similar to the associated samples and is known to be free of the analytes of interest.

4. Evaluation criteria and corrective action. While the goal is to have no detectable contaminants, each method blank shall be critically evaluated as to the nature of the interference and the effect on the analysis of each sample within the batch. The source of contamination shall be investigated and measures taken to minimize or eliminate the problem and affected samples reprocessed or data shall be appropriately qualified if:

a. The concentration of a targeted analyte in the blank is at or above the reporting limit as established by the test method or by regulation, and is greater than 1/10 of the amount measured in any sample.

b. The blank contamination otherwise affects the sample results as per the test method requirements or the individual project data quality objectives.

c. When a blank is determined to be contaminated, the cause shall be investigated and measures taken to minimize or eliminate the problem. Samples associated with a contaminated blank shall be evaluated as to the best corrective action for the samples (e.g. reprocessing or data qualifying codes). In all cases the corrective action shall be documented.

B. Positive control – method performance. Laboratory control sample (LCS).

1. Purpose. The LCS is used to evaluate the performance of the total analytical system, including all preparation and analysis steps. Results of the LCS are compared to established criteria and, if found to be outside of these criteria, indicates that the analytical system is "out of control." Any affected samples associated with an out of control LCS shall be reprocessed for re-analysis or the results reported with appropriate data qualifying codes.

2. Frequency. The LCS shall be analyzed at a minimum of one per preparation batch. Exceptions would be for those analytes for which no spiking solutions are available such as total suspended solids, total dissolved solids, total volatile solids, total solids, pH, color, odor, temperature, dissolved oxygen or turbidity. In those instances for which no separate preparation method is used (example: volatiles in water) the batch shall be defined as environmental samples that are analyzed together with the same method and personnel, using the same lots of reagents, not to exceed the analysis of 20 environmental samples.

3. Composition. The LCS is a quality system matrix, known to be free of analytes of interest, spiked with known and verified concentrations of analytes. NOTE: the matrix spike may be used in place of this control as long as the acceptance criteria are as stringent as for the LCS. Alternatively the LCS may consist of a media containing known and verified concentrations of analytes or as Certified Reference Material (CRM). All analyte concentrations shall be within the calibration range of the methods. The following shall be used in choosing components for the spike mixtures:

The components to be spiked shall be as specified by the mandated test method or other regulatory requirement or as requested by the client. In the absence of specified spiking components the laboratory shall spike per the following:

a. For those components that interfere with an accurate assessment such as spiking simultaneously with technical chlordane, toxaphene and PCBs, the spike should be chosen that represents the chemistries and elution patterns of the components to be reported.

b. For those test methods that have extremely long lists of analytes, a representative number may be chosen. The analytes selected should be representative of all analytes reported. The following criteria shall be used for determining the minimum number of analytes to be spiked. However, the laboratory shall insure that all targeted components are included in the spike mixture over a two-year period. For methods that include 1-10 targets, spike all components; for methods that include 11-20 targets, spike at least 10% or 80%, whichever is greater; and for methods with more than 20 targets, spike at least 16 components.

4. Evaluation criteria and corrective action.

a. The results of the individual batch LCS are calculated in percent recovery or other appropriate statistical technique that allows comparison to established acceptance criteria. The laboratory shall document the calculation.

b. The individual LCS is compared to the acceptance criteria as published in the mandated test method. Where there are no established criteria, the laboratory shall determine internal criteria and document the method used to establish the limits or utilize client specified assessment criteria.

c. A LCS that is determined to be within the criteria effectively establishes that the analytical system is in control and validates system performance for the samples in the associated batch. Samples analyzed along with a LCS determined to be "out of control" shall be considered suspect and the samples reprocessed and re-analyzed or the data reported with appropriate data qualifying codes.

5. If a large number of analytes are in the LCS, it becomes statistically likely that a few will be outside control limits. This may not indicate that the system is out of control, therefore corrective action may not be necessary. Upper and lower marginal exceedance (ME) limits can be established to determine when corrective action is necessary. A ME is defined as being beyond the LCS control limit (3 standard deviations), but within the ME limits. ME limits are between 3 and 4 standard deviations around the mean.

a. The number of allowable marginal exceedances is based on the number of analytes in the LCS. If more analytes exceed the LCS control limits than is allowed, or if any one analyte exceeds the ME limits, the LCS fails and corrective action is necessary. This marginal exceedance approach is relevant for methods with long lists of analytes. It will not apply to target analyte lists with fewer than 11 analytes.

b. The number of allowable marginal exceedances is as follows:

| Number of analytes in LCS | Number of analytes allowed in ME of the LCS control limit |
| --- | --- |
| Greater than 90 | Five |
| 71-90 | Four |
| 51-70 | Three |
| 31-50 | Two |
| 11-30 | One |
| Fewer than 11 | None |

c. Marginal exceedances shall be random. If the same analyte exceeds the LCS control limit repeatedly, it is an indication of a systemic problem. The source of the error shall be located and corrective action taken. Laboratories shall have a written procedure to monitor the application of marginal exceedance allowance to the LCS to ensure random behavior.

C. Sample specific controls - general.

1. The laboratory shall document procedures for determining the effect of the sample matrix on method performance. These procedures relate to the analyses of quality system matrix specific Quality Control (QC) samples and are designed as data quality indicators for a specific sample using the designated test method. These controls alone are not used to judge laboratory performance.

2. Examples of matrix specific QC include: Matrix Spike (MS); Matrix Spike Duplicate (MSD); sample duplicates; and surrogate spikes. The laboratory shall have procedures in place for tracking, managing, and handling matrix specific QC criteria including spiking appropriate components at appropriate concentrations, calculating recoveries and relative percent difference, evaluating and reporting results based on performance of the QC samples.

D. Sample specific controls - matrix spike and matrix spike duplicates.

1. Purpose. Matrix specific QC samples indicate the effect of the sample matrix on the precision and accuracy of the results generated using the selected method. The information from these controls is sample/matrix specific and would not normally be used to determine the validity of the entire batch.

2. Frequency. The frequency of the analysis of matrix specific samples shall be determined as part of a systematic planning process (e.g. Data Quality Objectives) or as specified by the test method.

3. Composition. The components to be spiked shall be as specified by the mandated test method. Any permit specified analytes, as specified by regulation or client requested analytes shall also be included. If there are no specified components, the laboratory shall spike per the following:

a. For those components that interfere with an accurate assessment such as spiking simultaneously with technical chlordane, toxaphene and PCBs, the spike should be chosen that represents the chemistries and elution patterns of the components to be reported.

b. For those test methods that have extremely long lists of analytes, a representative number may be chosen using the following criteria for choosing the number of analytes to be spiked. However, the laboratory shall insure that all targeted components are included in the spike mixture over a two-year period.

(1) For methods that include 1-10 targets, spike all components;

(2) For methods that include 11-20 targets, spike at least 10% or 80%, whichever is greater;

(3) For methods with more than 20 targets, spike at least 16 components.

4. Evaluation criteria and corrective action.

a. The results from matrix spike/matrix spike duplicate are primarily designed to assess the precision and accuracy of analytical results in a given matrix and are expressed as percent recovery (%R), relative percent difference (RPD), or other appropriate statistical technique that allows comparison to established acceptance criteria. The laboratory shall document the calculation for %R, RPD or other statistical treatment used.

b. The results are compared to the acceptance criteria as published in the mandated test method. Where there are no established criteria, the laboratory shall determine internal criteria and document the method used to establish the limits. For matrix spike results outside established criteria corrective action shall be documented or the data reported with appropriate data qualifying codes.

E. Sample specific controls - matrix duplicates.

1. Purpose. Matrix duplicates are defined as replicate aliquots of the same sample taken through the entire analytical procedure. The results from this analysis indicate the precision of the results for the specific sample using the selected method. The matrix duplicate provides a usable measure of precision only when target analytes are found in the sample chosen for duplication.

2. Frequency. The frequency of the analysis of matrix duplicates may be determined as part of a systematic planning process (e.g. Data Quality Objectives) or as specified by the mandated test method.

3. Composition. Matrix duplicates are performed on replicate aliquots of actual samples. The composition is usually not known.

4. Evaluation criteria and corrective action.

a. The results from matrix duplicates are primarily designed to assess the precision of analytical results in a given matrix and are expressed as relative percent difference (RPD) or another statistical treatment (e.g., absolute differences). The laboratory shall document the calculation for relative percent difference or other statistical treatments.

b. Results are compared to the acceptance criteria as published in the mandated test method. Where there are no established criteria, the laboratory shall determine internal criteria and document the method used to establish the limits. For matrix duplicates results outside established criteria corrective action shall be documented or the data reported with appropriate data qualifying codes.

F. Sample specific controls - surrogate spikes.

1. Purpose. Surrogates are used most often in organic chromatography test methods and are chosen to reflect the chemistries of the targeted components of the method. Added prior to sample preparation/extraction, they provide a measure of recovery for every sample matrix.

2. Frequency. Except where the matrix precludes its use or when not commercially available, surrogate compounds shall be added to all samples, standards, and blanks for all appropriate test methods.

3. Composition. Surrogate compounds are chosen to represent the various chemistries of the target analytes in the method or MQO. They are often specified by the mandated method and are deliberately chosen for their being unlikely to occur as an environmental contaminant. Often this is accomplished by using deuterated analogs of select compounds.

4. Evaluation criteria and corrective action. The results are compared to the acceptance criteria as published in the mandated test method. Where there are no established criteria, the laboratory should determine internal criteria and document the method used to establish the limits. Surrogates outside the acceptance criteria shall be evaluated for the effect indicated for the individual sample results.  Data quality objectives or other site-specific requirements may guide the appropriate corrective action.  Results reported from analyses with surrogate recoveries outside the acceptance criteria should include appropriate data qualifiers.

1VAC30-45-771. Chemical testing: limit of detection and limit of quantitation.

A. General. All procedures used shall be documented. Documentation shall include the quality system matrix type. All supporting data shall be retained.

B. Limit of detection (LOD). The laboratory shall utilize a test method that provides an LOD that is appropriate and relevant for the intended use of the data. An LOD is not required for a test method when test results are not reported outside of the calibration range. LODs shall be determined by the protocol in the mandated test method or applicable regulation. If the protocol for determining LODs is not specified, the selection of the procedure shall reflect instrument limitations and the intended application of the test method.

1. The LOD shall be initially determined for the compounds of interest in each test method in a quality system matrix in which there are no target analytes or interferences at a concentration that would impact the results. Alternatively the LOD shall be determined in the quality system matrix of interest (see definition of matrix).

2. LODs shall be determined each time there is a change in the test method that affects how the test is performed, or when a change in instrumentation occurs that affects the sensitivity of the analysis.

3. The laboratory shall have established procedures to relate LOD with LOQ.

4. The LOD shall be verified annually for each quality system matrix, method and analyte according to the procedure specified in 1VAC30-45-760 B 1.

C. Limit of quantitation (LOQ).

1. Any established LOQ shall be above the LOD.

2. The LOQ shall be verified annually for each quality system matrix, method and analyte according to the procedure specified in 1VAC30-45-760 B 2. Alternatively, the annual LOQ verification is not required if the LOD is reevaluated or verified according to subdivision B 4 of this section.

1VAC30-45-772. Chemical testing: data reduction.

The procedures for data reduction, such as use of linear regression, shall be documented.

1VAC30-45-773. Chemical testing: quality of standards and reagents.

A. The source of standards shall comply with 1VAC30-45-740 C.

B. Reagent quality, water quality and checks.

1. Reagents. In methods where the purity of reagents is not specified, analytical reagent grade shall be used. Reagents of lesser purity than those specified by the test method shall not be used. The labels on the container should be checked to verify that the purity of the reagents meets the requirements of the particular test method. Such information shall be documented.

2. Water. The quality of water sources shall be monitored and documented and shall meet method specified requirements.

3. The laboratory will verify the concentration of titrants in accordance with written laboratory procedures.

1VAC30-45-774. Chemical testing: selectivity.

A. The laboratory shall evaluate selectivity by following the checks established within the method, which may include mass spectral tuning, second column confirmation, ICP inter-element interference checks, chromatography retention time windows, sample blanks, spectrochemical absorption or fluorescence profiles, co-precipitation evaluations, and electrode response factors.

B. A confirmation shall be performed to verify the compound identification when positive results are detected on a sample from a location that has not been previously tested by the laboratory. Such confirmations shall be performed on organic tests such as pesticides, herbicides, or acid extractable or when recommended by the analytical test method except when the analysis involves the use of a mass spectrometer. Confirmation is required unless stipulated in writing by the client. All confirmation shall be documented.

C. The laboratory shall document acceptance criteria for mass spectral tuning.

1VAC30-45-775. Chemical testing: constant and consistent test conditions.

A. The laboratory shall assure that the test instruments consistently operate within the specifications required of the application for which the equipment is used.

B. Glassware cleaning. Glassware shall be cleaned to meet the sensitivity of the test method.

C. Any cleaning and storage procedures that are not specified by the test method shall be documented in laboratory records and SOPs.

1VAC30-45-780. Toxicity testing: general.

These standards apply to laboratories measuring the toxicity and/or bioaccumulation of contaminants in effluents (aquatic toxicity), receiving waters, sediments, elutriates, leachates and soils. In addition to the essential quality control standards set out in 1VAC30-45-781 through 1VAC30-45-788, some methods may have additional or other requirements based on factors such as the type of quality system matrix evaluated.

1VAC30-45-781. Toxicity testing: positive and negative controls.

A. Positive control. Reference toxicant tests demonstrate a laboratory's ability to obtain consistent results with the test method and evaluate the overall health and sensitivity of test organisms over time.

1. The laboratory shall demonstrate its ability to obtain consistent results with standard reference toxicants (SRT) and complete an initial Demonstration of Capability (DOC) in order to attain accreditation in toxicity testing methods.

a. An initial DOC shall consist of five or more acceptable SRT tests for each test method, species and endpoint with different batches of organisms. Appropriate negative controls (water, sediment, or soil) shall be tested at the frequency and duration specified in the test method. Initial DOCs shall be prepared in accordance with the requirements of  1VAC30-45-730 F.

b. Initial DOC is established by maintenance of SRT test results on control charts. A laboratory shall record the control performance and statistical endpoints (such as NOEC or ECp) for each method species and endpoint on control charts. Initial DOC is established where 95% of the test results required in subdivision A 1 a of this section fall within the control limits established in accordance with subdivision A 1 c of this section and meet test acceptability criteria (TAC). The laboratory shall evaluate precision (i.e., coefficient of variation (CV)) or sensitivity  (i.e., statistical minimum significant difference (SMSD) measures; see subdivision A 1 d of this section) for these tests against method-specific or, lacking the former, laboratory-derived criteria to determine validity of the initial DOC.

c. For endpoints that are point estimates (ICp, ECp), control charts are constructed by plotting the cumulative mean and the control limits that consist of the upper and lower 95% confidence limits (+/- 2 standard deviations). In case of highly variable point estimates that exceed method-specific criteria, the control chart limits are adjusted accordingly. For endpoints from hypothesis tests (NOEC, NOAEC), the values are plotted directly and the control limits consist of one concentration interval above and below the concentration representing the central tendency (i.e., the mode).

d. For endpoints that are point estimates, the cumulative mean CV is calculated and for endpoints from hypothesis tests, the SMSD is calculated. These values are maintained on a control chart.

2. Ongoing laboratory performance shall be demonstrated by routine SRT testing for each test method and species and endpoint in accordance with the minimum frequency requirements specified in subdivision A 3 of this section.

a. Intralaboratory precision is determined on an ongoing basis through the use of control charts as established in subdivision A 1 b of this section. The control charts shall be plotted as point estimate values, such as EC25 for chronic tests and LC50 for acute tests, or as appropriate hypothesis test values, such as the NOEC or NOAEC, over time within a laboratory.

b. After initial laboratory DOC is determined, the control limits and CV for an individual test method, endpoints and species shall be adjusted as additional test results are obtained. After 20 data points are collected for a test method and species, the control chart is maintained using only the last 20 data points, i.e. each successive mean value and control limit is calculated using only the last 20 values.

c. Control chart limits are expected to be exceeded occasionally regardless of how well a laboratory performs. Acceptance limits for point estimates (ICp, ECp) that are based on 95% confidence limits should theoretically be exceeded for one in 20 tests. Depending on the dilution factor and test sensitivity, control charts based on hypothesis test values (NOEC, NOAEC) may be expected to be exceeded on a similar frequency. Test results that fall outside of control chart limits at a frequency of 5.0% or less, or that fall just outside control chart limits (especially in the case of highly proficient laboratories that may develop relatively narrow acceptance limits over time), are not rejected de facto. Such data are evaluated in comparison with control chart characteristics including the width of the acceptance limits and the degree of departure of the value from acceptance limits.

d. Consistent with the test methods used, laboratories shall develop acceptance/rejection policies for SRT data that consider the source of test organisms, the direction of the deviation, test dilution factor, test sensitivity (for hypothesis test values), testing frequency, out-of-control test frequency, relative width of acceptance limits, inter-test CV, and degree of difference between test results and acceptance limits.

e. In the case of reference toxicant data that fails to meet control chart acceptance criteria, the test data are examined for defects, corrective action taken, and the test repeated if necessary, using a different batch of organisms or the data is qualified.

3. The frequency of ongoing laboratory reference toxicant testing shall be as follows unless the method specifically requires less frequent SRT tests (e.g., sediment tests):

a. For test methods conducted at a frequency of monthly or greater, SRT tests shall be conducted at an ongoing frequency of monthly.

b. For test methods and species commonly used in the laboratory, but that are tested at a frequency of less than monthly, SRT tests shall be conducted concurrently with the environmental test.

c. If the test organisms are obtained from an outside source the sensitivity of each batch of organisms received from a supplier shall be determined via a concurrent SRT test unless the supplier can provide control chart data for the last five SRT tests using the same SRT and test conditions. Supplied SRT data may not be older than six months.

d. The DOC for an analyst shall be consistent with 1VAC30-45-220 B but the frequency need not exceed the method-specified requirements and subdivision A 3 a and A 3 b of this section.

4. These standards do not currently specify a particular reference toxicant and dilution series. If the  permitting authority  identifies a reference toxicant or dilution series for a particular test, the laboratory shall follow the specified requirements. All reference toxicant tests conducted for a given test method and species shall use the same reference toxicant, test concentrations, dilution water and data analysis methods. A dilution factor of 0.5x or greater shall be used for both acute and chronic tests.

5. The reference toxicant tests shall be conducted following the same procedures as the environmental toxicity tests for which the precision is being evaluated, unless otherwise specified in the test method (for example, 10-day sediment tests employ 96-h water-only reference toxicant tests). The test duration, laboratory dilution water, feeding, organism age, range and density, test volumes, renewal frequency, water quality measurements, and the number of test concentrations, replicates and organisms per replicate shall be the same as specified for the environmental toxicity test.

B. Negative control: control, brine control, control sediment, control soil or dilution water.

1. The standards for the use, type and frequency of testing of negative controls are specified by the test methods and by permit or regulation and shall be followed. A negative control is included with each test to evaluate test performance and the health and sensitivity of the specific batch of organisms.

2. Appropriate additional negative controls shall be included when sample adjustments (for example, addition of thiosulfate for dechlorination) or solvent carriers are used in the test.

3. Test acceptability criteria (TAC). The test acceptability criteria specified in the test method shall be achieved for both the reference toxicant and the effluent or environmental sample toxicity test. The criteria shall be calculated and shall meet the method specified requirements for performing toxicity tests.

1VAC30-45-782. Toxicity testing: variability and/or reproducibility.

Intralaboratory precision shall be determined on an ongoing basis through the use of further reference toxicant tests and related control charts as described in 1VAC30-45-840 A.

1VAC30-45-783. Toxicity testing: accuracy.

This principle is not applicable to toxicity testing.

1VAC30-45-784. Toxicity testing: test sensitivity.

A. The statistical minimum significant difference (SMSD) shall be calculated according to the formula specified by the test method and reported with the test results.

B. Point estimates: (LCp, ICp, or ECp) Confidence intervals shall be reported as a measure of the precision around the point estimate value, when the calculation is possible.

C. The SMSD shall be calculated and reported for only hypothesis test values, such as the NOEC or NOAEC.

1VAC30-45-785. Toxicity testing: selection of appropriate statistical analysis methods.

A. If required, methods of data analysis and endpoints are specified by language in the regulation, permit or the test method.

B. Dose response curves. The data shall be plotted in the form of a curve relating the dose of the chemical or concentration of sample to cumulative percentage of test organisms demonstrating a response such as death. Evaluation criteria shall be established for interpretation of concentration or dose response curves.

1VAC30-45-786. Toxicity testing: selection and use of reagents and standards.

A. The grade of all reagents used in toxicity tests is specified in the test method except the reference standard. All reference standards shall be prepared from chemicals that are analytical reagent grade or better. The preparation of all standards and reference toxicants shall be documented.

B. All standards and reagents associated with chemical measurements, such as dissolved oxygen, pH or specific conductance, shall comply with the standards outlined in 1VAC30-45-740 D 1 d.

C. Only reagent-grade water collected from distillation or deionization units is used to prepare reagents.

1VAC30-45-787. Toxicity testing: selectivity.

The permit or regulation specifies the selectivity of the test.

1VAC30-45-788. Toxicity testing: constant and consistent test conditions.

A. If closed refrigerator-sized incubators are used, culturing and testing of organisms shall be separated to avoid cross-contamination.

B. Laboratory space shall be adequate for the types and numbers of tests performed. The building shall provide adequate cooling, heating and illumination for conducting testing and culturing; hot and cold running water shall be available for cleaning equipment.

C. Air used for aeration of test solutions, dilution waters and cultures shall be free of oil and fumes.

D. The laboratory or a contracted outside expert shall positively identify test organisms to species on an annual basis. The taxonomic reference (citation and page(s)) and the names(s) of the taxonomic expert(s) shall be kept on file at the laboratory. When organisms are obtained from an outside source, the supplier shall provide this same information.

E. Instruments used for routine support measurements of chemical and physical parameters such as pH, DO, conductivity, salinity, alkalinity, hardness, chlorine, ammonia, and weight shall be calibrated, and/or standardized per manufacturer’s instructions. As these are support measurements, only the calibration and verification requirements specified at 1VAC30-45-740 D 1 apply. All measurements and calibrations shall be documented.

F. Test temperature shall be maintained as specified for the test method. Temperature control equipment shall be adequate to maintain the required test temperature(s). The average daily temperature of the test solutions shall be maintained within the method-specified range. The minimum frequency of measurement shall be once per 24-hour period. The test temperature for continuous-flow toxicity tests shall be recorded and monitored continuously. Where electronic data loggers are used, temperature shall be monitored at a frequency sufficient to capture temporal variations of the environmental control system.

G. Reagent-grade water, prepared by any combination of distillation, reverse osmosis, ion exchange, activated carbon and particle filtration, shall meet the method specified requirements.

H. The quality of the standard dilution water used for testing or culturing shall be sufficient to allow satisfactory survival, growth and reproduction of the test species as demonstrated by routine reference toxicant tests and negative control performance. Water used for culturing and testing shall be analyzed for toxic metals and organics whenever the minimum acceptability criteria for control survival, growth or reproduction are not met and no other cause, such as contaminated glassware or poor stock, can be identified. It is recognized that the analyte lists of some methods manuals may not include all potential toxicants, are based on estimates of chemical toxicity available at the time of publication and may specify detection limits that are not achievable in all matrices. However, for those analytes not listed, or for which the measured concentration or limit of detection is greater than the method-specified limit, the laboratory shall demonstrate that the analyte at the measured concentration or reported limit of detection does not exceed one-tenth of the expected chronic value for the most sensitive species tested and/or cultured. The expected chronic value is based on professional judgment and the best available scientific data. The "USEPA Ambient Water Quality Criteria Documents" and the EPA AQUIRE database provide guidance and data on acceptability and toxicity of individual metals and organic compounds.

I. The quality of the food used for testing or culturing shall be sufficient to allow satisfactory survival, growth and reproduction of the test species as demonstrated by routine reference toxicant tests and negative control performance. The laboratory shall have written procedures for the evaluation of food acceptance.

J. A subset of organisms used in bioaccumulation tests shall be analyzed at the start of the test (baseline) for the target compounds to be measured in the bioaccumulation tests.

K. Test chamber size and test solution volume shall be as specified in the test method. All test chambers used in a test shall be identical.

L. Test organisms shall be fed the quantity and type food or nutrients specified in the test method. They shall also be fed at the intervals specified in the test methods.

M. All organisms in a test shall be from the same source. Where available certified seeds are used for soil tests.

N. All organisms used in tests, or used as broodstock to produce neonate test organisms (for example cladocerans and larval fish), shall appear healthy, show no signs of stress or disease and exhibit acceptable survival (90% or greater) during the 24-hour period immediately preceding use in tests.

O. All materials used for test chambers, culture tanks, tubing, etc., and coming in contact with test samples, solutions, control water, sediment or soil or food shall be nontoxic and cleaned as described in the test methods. Materials shall not reduce or add to sample toxicity. Appropriate materials for use in toxicity testing and culturing are described in the referenced manuals.

P. Light intensity shall be maintained as specified in the methods manuals. Measurements shall be made and recorded on a yearly basis. Photoperiod shall be maintained as specified in the test methods and shall be documented at least quarterly. For algal and plant tests, the light intensity shall be measured and recorded at the start of each test.

Q. The testing laboratory shall document the health and culturing conditions of all organisms used for testing. Such documentation shall include culture conditions (e.g., salinity, hardness, temperature, pH) and observations of any stress, disease or mortality. When organisms are obtained from an outside source, the laboratory shall obtain written documentation of these water quality parameters and biological observations for each lot of organism received. These observations shall adequately address the 24-hour time period referenced in subsection N of this section. The laboratory shall also record each of these observations and water quality parameters upon the arrival of the organisms at the testing laboratory.

R. Age and the age range of the test organisms shall be as specified in the test method. Supporting information, such as hatch dates and times, times of brood releases and metrics (for example, chironomid head capsule width) shall be documented.

S. The maximum holding time of effluents (elapsed time from sample collection to first use in a test) shall not exceed 36 hours; samples may be used for renewal up to 72 hours after first use except as prescribed by the method and approved by the regulatory agency having authority for program oversight.

T. All samples shall be chilled to 0 to 6°C during or immediately after collection except as prescribed by the method.

U. Organisms used in a given test shall be from the same batch.

V. All tests shall have the minimum number of replicates per treatment as prescribed by the method.

W. The control population of Ceriodaphnia in chronic effluent or receiving water tests shall contain no more than 20% males.

X. The culturing of C. dubia shall be adequate such that blocking by parentage can be established.

Y. Dissolved oxygen and pH in aquatic tests shall be within acceptable range at test initiation and aeration (minimal) is provided to tests if, and only if, acceptable dissolved oxygen concentrations cannot be otherwise maintained or if specified by the test method.

Z. Test soils or sediments shall be within the geochemical tolerance range of the test organism.

AA. An individual test may be conditionally acceptable if temperature, dissolved oxygen, pH and other specified conditions fall outside specifications, depending on the degree of the departure and the objectives of the tests (see test conditions and test acceptability criteria specified for each test method).

1VAC30-45-789. (Reserved.)

1VAC30-45-790. Microbiology testing: general.

These standards apply to laboratories undertaking microbiological analysis of environmental samples. Microbiological testing refers to and includes the detection, isolation, enumeration, or identification of microorganisms and/or their metabolites, or determination of the presence or absence of growth in materials and media.

1VAC30-45-791. Microbiology testing: sterility checks and blanks, positive and negative controls.

A. Sterility checks and blanks. The laboratory shall demonstrate that the filtration equipment and filters, sample containers, media and reagents have not been contaminated through improper handling or preparation, inadequate sterilization, or environmental exposure.

1. A sterility blank shall be analyzed for each lot of pre-prepared, ready-to-use medium (including chromofluorogenic reagent) and for each batch of medium prepared in the laboratory. This shall be done prior to first use of the medium.

2. For filtration technique, the laboratory shall conduct one beginning and one ending sterility check for each laboratory sterilized filtration unit used in a filtration series. The filtration series may include single or multiple filtration units, which have been sterilized prior to beginning the series. For presterilized single use funnels a sterility check shall be performed on one funnel per lot. The filtration series is considered ended when more than 30 minutes elapses between successive filtrations. During a filtration series, filter funnels shall be rinsed with three 20-30 ml portions of sterile rinse water after each sample filtration. In addition, laboratories shall insert a sterility blank after every 10 samples or sanitize filtration units by UV light after each sample filtration.

3. For pour plate technique, sterility blanks of the medium shall be made by pouring, at a minimum, one uninoculated plate for each lot of pre-prepared, ready-to-use media and for each batch of medium prepared in the laboratory.

4. Sterility checks on sample containers shall be performed on at least one container for each lot of purchased, presterilized containers. For containers prepared and sterilized in the laboratory, a sterility check shall be performed on one container per sterilized batch with nonselective growth media.

5. A sterility blank shall be performed on each batch of dilution water prepared in the laboratory and on each batch of pre-prepared, ready-to-use dilution water with nonselective growth media.

6. At least one filter from each new lot of membrane filters shall be checked for sterility with nonselective growth media.

B. Positive controls.

1. Positive culture controls demonstrate that the medium can support the growth of the target organism(s), and that the medium produces the specified or expected reaction to the target organism(s).

2. Each preprepared, ready-to-use lot of medium (including chromofluorogenic reagent) and each batch of medium prepared in the laboratory shall be tested  and demonstrate a known positive response. This shall be done prior to first use of the medium.

C. Negative controls.  The provisions of this subsection shall not apply to wastewater treatment plants.

1. Negative culture controls demonstrate that the medium does not support the growth of non-target organisms or does not demonstrate the typical positive reaction of the target organism(s).

2. Each pre-prepared, ready-to-use lot of selective medium (including chromofluorogenic reagent) and each batch of selective medium prepared in the laboratory shall be analyzed with one or more known negative culture controls, i.e. nontarget organisms, as appropriate to the method. This shall be done prior to first use of the medium.

1VAC30-45-792. Microbiology testing: test variability and reproducibility .

For test methods that specify colony counts such as membrane filter or plated media, duplicate counts shall be performed monthly on one positive sample, for each month that the test is performed. If the lab has two or more analysts, each analyst shall count typical colonies on the same plate. Counts shall be within 10% difference to be acceptable. In a laboratory with only one microbiology analyst, the analyst shall count the same plate twice, with no more than 5.0% difference between the counts.

1VAC30-45-793. Microbiology testing: method evaluation.

A. Laboratories are required to demonstrate proficiency with the test method prior to first use. This shall be achieved by comparison to a method already approved for use in the laboratory, or by analyzing a minimum of 10 spiked samples whose quality system matrix is representative of those normally submitted to the laboratory, or by analyzing and passing one proficiency test series provided by an approved proficiency sample provider. The laboratory shall maintain this documentation as long as the method is in use and for at least five years past the date of last use.

B. Laboratories shall participate in the proficiency test programs required by Article 3 (1VAC30-45-500 et seq.). The results of these analyses shall be used to evaluate the ability of the laboratory to produce acceptable data.

1VAC30-45-794. Microbiology testing: test performance.

A. All growth and recovery media shall be checked to assure that the target organism(s) respond in an acceptable and predictable manner (see 1VAC30-45-791 B).

B. To ensure that analysis results are accurate, target organism identity shall be verified as specified in the method, e.g., by use of the completed test, or by use of secondary verification tests such as a catalase test.

1VAC30-45-795. Microbiology testing: data reduction.

The calculations, data reduction and statistical interpretations specified by each test method shall be followed.

1VAC30-45-796. Microbiology testing: quality of standards, reagents and media.

A. The laboratory shall ensure that the quality of the reagents and media used is appropriate for the test concerned.

B. Culture media may be prepared from commercial dehydrated powders or may be purchased ready to use. The laboratory may prepare media from basic ingredients when commercial media are not available or when it can be demonstrated that commercial media do not provide adequate results. Media prepared by the laboratory from basic ingredients shall be tested for performance (e.g., for selectivity, sensitivity, sterility, growth promotion, growth inhibition) prior to first use. Detailed testing criteria information shall be defined in either the laboratory’s test methods, SOPs, quality manual, or similar documentation.

C. Reagents, commercial dehydrated powders and media shall be used within the shelf-life of the product and shall be documented according to 1VAC30-45-730 J.

D. Distilled water, deionized water or reverse osmosis produced water free from bactericidal and inhibitory substances shall be used in the preparation of media, solutions and buffers. The quality of the water shall be monitored for chlorine residual, specific conductance, and heterotrophic bacteria plate count monthly (when in use), when maintenance is performed on the water treatment system, or at startup after a period of disuse longer than one month.

E. Analysis for metals and the Bacteriological Water Quality Test (to determine presence of toxic agents or growth promoting substances) shall be performed annually. Results of these analyses shall meet the specifications of the required method and records of analyses shall be maintained for three years. (An exception to performing the Bacteriological Water Quality Test shall be given to laboratories that can supply documentation to show that their water source meets the criteria, as specified by the method, for Type I or Type II reagent water.)

F. Media, solutions and reagents shall be prepared, used and stored according to a documented procedure following the manufacturer’s instructions or the test method. Documentation for media prepared in the laboratory shall include date of preparation, preparer’s initials, type and amount of media prepared, manufacturer and lot number, final pH of the media, and expiration date. Documentation for media purchased pre-prepared, ready to use shall include manufacturer, lot number, type and amount of media received, date of receipt, expiration date of the media, and pH of the media.

1VAC30-45-797. Microbiology testing: selectivity.

In order to ensure identity and traceability, reference cultures used for positive and negative controls shall be obtained from a recognized national collection, organization, or manufacturer. Microorganisms may be single use preparations or cultures maintained by documented procedures that demonstrate the continued purity and viability of the organism.

1. Reference cultures may be revived (if freeze-dried) or transferred from slants and subcultured once to provide reference stocks. The reference stocks shall be preserved by a technique that maintains the characteristics of the strains. Reference stocks shall be used to prepare working stocks for routine work. If reference stocks have been thawed, they shall not be refrozen and reused.

2. Working stocks shall not be sequentially cultured more than five times and shall not be subcultured to replace reference stocks.

1VAC30-45-798. Microbiology testing: constant and consistent test conditions.

A. Laboratory facilities. Floors and work surfaces shall be nonabsorbent and easy to clean and disinfect. Work surfaces shall be adequately sealed. Laboratories shall provide sufficient storage space, and shall be clean and free from dust accumulation. Plants, food, and drink shall be prohibited from the laboratory work area.

B. Laboratory equipment.

1. Temperature measuring devices. Temperature measuring devices such as liquid-in-glass thermometers, thermocouples, and platinum resistance thermometers used in incubators, autoclaves and other equipment shall be the appropriate quality to meet specification(s) in the test method. The graduation of the temperature measuring devices shall be appropriate for the required accuracy of measurement and they shall be calibrated to national or international standards for temperature (see 1VAC30-45-740 C). Calibration shall be done at least annually.

2. Autoclaves.

a. The performance of each autoclave shall be initially evaluated by establishing its functional properties and performance, for example, heat distribution characteristics with respect to typical uses. Autoclaves shall meet specified temperature tolerances. Pressure cookers shall not be used for sterilization of growth media.

b. Demonstration of sterilization temperature shall be provided by use of continuous temperature recording device or by use of a maximum registering thermometer with every cycle. Appropriate biological indicators shall be used once per month to determine effective sterilization. Temperature sensitive tape shall be used with the contents of each autoclave run to indicate that the autoclave contents have been processed.

c. Records of autoclave operations shall be maintained for every cycle. Records shall include date, contents, maximum temperature reached, pressure, time in sterilization mode, total run time (may be recorded as time in and time out) and analyst’s initials.

d. Autoclave maintenance, either internally or by service contract, shall be performed annually and shall include a pressure check and calibration of temperature device. Records of the maintenance shall be maintained in equipment logs.

e. The autoclave mechanical timing device shall be checked quarterly against a stopwatch and the actual time elapsed documented.

3. Volumetric equipment. Volumetric equipment shall be calibrated as follows:

a. Equipment with movable parts such as automatic dispensers, dispensers/diluters, and mechanical hand pipettes shall be verified for accuracy quarterly.

b. Equipment such as filter funnels, bottles, nonclass A glassware, and other marked containers shall be calibrated once per lot prior to first use.

c. The volume of the disposable volumetric equipment such as sample bottles and disposable pipettes shall be checked once per lot.

4. UV instruments. UV instruments used for sanitization shall be tested quarterly for effectiveness with an appropriate UV light meter or by plate count agar spread plates. Replace bulbs if output is less than 70% of original for light tests or if count reduction is less than 99% for a plate containing 200 to 300 organisms.

5. Conductivity meters, oxygen meters, pH meters, hygrometers, and other similar measurement instruments shall be calibrated according to the method specified requirements (see 1VAC30-45-740 D 1 d).

6. Incubators, water baths, and ovens.

a. The stability and uniformity of temperature distribution and time required after test sample addition to reestablish equilibrium conditions in incubators and water baths shall be established. Temperature of incubators and water baths shall be documented twice daily, at least four hours apart, on each day of use.

b. Ovens used for sterilization shall be checked for sterilization effectiveness monthly with appropriate biological indicators. Records shall be maintained for each cycle that include date, cycle time, temperature, contents and analyst’s initials.

7. Labware (glassware and plasticware).

a. The laboratory shall have a documented procedure for washing labware, if applicable. Detergents designed for laboratory use shall be used.

b. Glassware shall be made of borosilicate or other noncorrosive material, free of chips and cracks, and shall have readable measurement marks.

c. Labware that is washed and reused shall be tested for possible presence of residues that may inhibit or promote growth of microorganisms by performing the Inhibitory Residue Test annually, and each time the lab changes the lot of detergent or washing procedures.

d. Washed labware shall be tested at least once daily, each day of washing, for possible acid or alkaline residue by testing at least one piece of labware with a suitable pH indicator such as bromothymol blue. Records of tests shall be maintained.

1VAC30-45-800. Radiochemical testing: general.

These standards apply to laboratories undertaking the examination of environmental samples by radiochemical analysis. These procedures for radiochemical analysis may involve some form of chemical separation followed by detection of the radioactive decay of analyte (or indicative daughters) and tracer isotopes where used. For the purpose of these standards, procedures for the determination of radioactive isotopes by mass spectrometry (e.g., ICP-MS or TIMS) or optical (e.g., KPA) techniques are not addressed herein.

1VAC30-45-801. Radiochemical testing: negative and positive controls.

A. Negative controls.

1. Method blank shall be performed at a frequency of one per preparation batch. The results of this analysis shall be one of the quality control measures to be used to assess the batch. The method blank result shall be assessed against the specific acceptance criteria specified in the laboratory method manual. When the specified method blank acceptance criteria is not met, the specified corrective action and contingencies shall be followed and results reported with appropriate data qualifying codes. The occurrence of a failed method blank acceptance criteria and the actions taken shall be noted in the laboratory report.

2. In the case of gamma spectrometry, generally a nondestructive analysis, a method blank shall be prepared using a calibrated counting geometry similar to that used for the samples. The container of the appropriate geometry can be empty or filled to similar volume to partially simulate gamma attenuation due to a sample matrix.

3. There shall be no subtraction of the required method blank result from the sample results in the associated preparation or analytical batch unless permitted by method or program. This does not preclude the application of any correction factor (e.g. instrument background, analyte presence in tracer, reagent impurities, peak overlap, etc.) to all analyzed samples, both program/project submitted and internal quality control samples. However, these correction factors shall not depend on the required method blank result in the associated analytical batch.

4. The method blank sample shall be prepared with similar aliquot size to that of the routine samples for analysis and the method blank result and acceptance criteria shall be calculated in a manner that compensates for sample results based upon differing aliquot size.

B. Positive controls.

1. Laboratory control samples shall be performed at a frequency of one per preparation batch. The results of this analysis shall be one of the quality control measures to be used to assess the batch. The laboratory control sample result shall be assessed against the specific acceptance criteria specified in the laboratory method manual. When the specified laboratory control sample acceptance criteria is not met the specified corrective action and contingencies shall be followed. The occurrence of a failed laboratory control sample acceptance criteria and the actions taken shall be noted in the laboratory report.

2. Matrix spike shall be performed at a frequency of one per preparation batch for those methods that include a chemical separation process without the use of an internal standard or carrier, and where there is sufficient sample to do so. Although gross alpha, gross beta and tritium measurements do not involve a chemical separation process, matrix spikes shall be performed for these analyses on aqueous samples. The results of this analysis shall be one of the quality control measures to be used to assess the batch. The matrix spike result shall be assessed against the specific acceptance criteria specified in the laboratory method manual. When the specified matrix spike acceptance criteria is not met, the specified corrective action and contingencies shall be followed. The occurrence of a failed matrix spike acceptance criteria and the actions taken shall be noted in the laboratory report. The lack of sufficient sample aliquot size to perform a matrix spike shall be noted in the laboratory report.

3. The activity of the laboratory control sample shall (i) be at least five times the limit of detection and (ii) at a level comparable to that of routine samples when such information is available if the sample activities are expected to exceed five times the limit of detection.

4. The activity of the matrix spike analytes(s) shall be greater than five times the limit of detection.

5. The laboratory standards used to prepare the laboratory control sample and matrix spike shall be from a source independent of the laboratory standards used for instrument calibration and shall meet the requirements for reference standards provided in 1VAC30-45-807 A.

6. The matrix spike shall be prepared by adding a known activity of target analyte after subsampling if required but before any chemical treatment (e.g., chemical digestion, dissolution, separation, etc.). Where a radiochemical method, other than gamma spectroscopy, has more than one reportable analyte isotope (e.g., plutonium, Pu 238 and Pu 239, using alpha spectrometry), only one of the analyte isotopes need be included in the laboratory control or matrix spike sample at the indicated activity level. However, where more than one analyte isotope is present above the specified limit of detection, each shall be assessed against the specified acceptance criteria.

7. Where gamma spectrometry is used to identify and quantitate more than one analyte isotope, the laboratory control sample shall contain isotopes that represent the low (e.g., americium-241), medium (e.g., cesium-137) and high (e.g., cobalt-60) energy range of the analyzed gamma spectra. As indicated by these examples the isotopes need not exactly bracket the calibrated energy range or the range over which isotopes are identified and quantitated.

8. The laboratory control sample shall be prepared with similar aliquot size to that of the routine samples for analyses.

C. Other controls.

1. Tracer. For those methods that utilize a tracer (i.e., internal standard) each sample result shall have an associated tracer recovery calculated and reported. The tracer shall be added to the sample after subsampling if required but before any chemical treatment (e.g., chemical digestion, dissolution, separation, etc.) unless otherwise specified by the method. The tracer recovery for each sample result shall be one of the quality control measures to be used to assess the associated sample result acceptance. The tracer recovery shall be assessed against the specific acceptance criteria specified in the laboratory method manual. When the specified tracer recovery acceptance criteria is not met the specified corrective action and contingencies shall be followed. The occurrence of a failed tracer recovery acceptance criteria and the actions taken shall be noted in the laboratory report.

2. Carrier. For those methods that utilize a carrier for recovery determination, each sample shall have an associated carrier recovery calculated and reported. The carrier shall be added to the sample after subsampling if required but before any chemical treatment (e.g., chemical digestion, dissolution, separation, etc.) unless otherwise specified by the method. The carrier recovery for each sample shall be one of the quality control measures to be used to assess the associated sample result acceptance. The carrier recovery shall be assessed against the specific acceptance criteria specified in the laboratory method manual. When the specified carrier recovery acceptance criteria is not met the specified corrective action and contingencies shall be followed. The occurrence of a failed carrier recovery acceptance criteria and the actions taken shall be noted in the laboratory report.

1VAC30-45-802. Radiochemical testing: analytical variability/reproducibility.

A. Replicate shall be performed at a frequency of one per preparation batch where there is sufficient sample to do so. The results of this analysis shall be one of the quality control measures to be used to assess batch acceptance. The replicate result shall be assessed against the specific acceptance criteria specified in the laboratory method manual. When the specified replicate acceptance criteria is not met the specified corrective action and contingencies shall be followed. The occurrence of a failed replicate acceptance criteria and the actions taken shall be noted in the laboratory report.

B. For low level samples (less than approximately three times the limit of detection) the laboratory may analyze duplicate laboratory control samples or a replicate matrix spike (matrix spike and a matrix spike duplicate) to determine reproducibility within a preparation batch.

1VAC30-45-803. Radiochemical testing: method evaluation.

In order to ensure the accuracy of the reported result, the following procedures shall be in place:

1. Initial demonstration of capability shall be performed initially (prior to the analysis of any samples) and with a significant change in instrument type (e.g., different detection technique), personnel or method.

2. Proficiency test samples. The laboratory shall use the results of such analysis to evaluate its ability to produce accurate data.

1VAC30-45-804. Radiochemical testing: radiation measurement instrumentation.

A. General. Because of the stability and response nature of modern radiation measurement instrumentation, it is not typically necessary to verify calibrate these systems each day of use. However, verification of calibration is required as outlined in subsection B of this section. This section addresses those practices that are necessary for proper calibration and those requirements of 1VAC30-45-740 D (instrument calibrations) that are not applicable to some types of radiation measurement instrumentation.

B. Instrument calibration.

1. Given that activity detection efficiency is independent of sample activity at all but extreme activity levels, the requirements of 1VAC30-45-740 D 2 b (7) are not applicable to radiochemical method calibrations except mass attenuation in gas-proportional counting and sample quench in liquid scintillation counting. Radiation measurement instruments are subject to calibration prior to initial use, when the instrument is placed back in service after malfunctioning and the instrument’s response has changed as determined by a performance check or when the instrument’s response exceeds predetermined acceptance criteria for the instrument quality control.

2. Instrument calibration shall be performed with reference standards as defined in 1VAC30-45-807 A. The standards shall have the same general characteristics (i.e., geometry, homogeneity, density, etc.) as the associated samples.

3. The frequency of calibration shall be addressed in the laboratory method manual if not specified in the method. A specific frequency (e.g., monthly) or observations from the associated control or tolerance chart, as the basis for calibration shall be specified.

C. Continuing instrument calibration verification (performance checks). Performance checks shall be performed using appropriate check sources and monitored with control charts or tolerance charts to ensure that the instrument is operating properly and that the detector response has not significantly changed and, therefore, the instrument calibration has not changed. The same check source used in the preparation of the tolerance chart or control chart at the time of calibration shall be used in the calibration verification of the instrument. The check sources shall provide adequate counting statistics for a relatively short count time and the source should be sealed or encapsulated to prevent loss of activity and contamination of the instrument and laboratory personnel.

1. For gamma spectroscopy systems, the performance checks for efficiency and energy calibration shall be performed on a day-of-use basis along with performance checks on peak resolution.

2. For alpha spectroscopy systems, the performance check for energy calibration shall be performed on a weekly basis and the performance check for counting efficiency shall be performed on at least a monthly basis.

3. For gas-proportional and liquid scintillation counters, the performance check for counting efficiency shall be performed on a day-of-use basis. For batches of samples that uninterruptedly count for more than a day a performance check can be performed at the beginning and end of the batch as long as this time interval is no greater than one week. Verification of instrument calibration does not directly verify secondary calibrations, e.g., the mass efficiency curve or the quench curve.

4. For scintillation counters the calibration verification for counting efficiency shall be performed on a day of use basis.

D. Background measurement. Background measurements shall be made on a regular basis and monitored using control charts or tolerance charts to ensure that a laboratory maintains its capability to meet required data quality objectives. These values may be subtracted from the total measured activity in the determination of the sample activity.

1. For gamma spectroscopy systems, background measurements shall be performed on at least a monthly basis.

2. For alpha spectroscopy systems, background measurements shall be performed on at least a monthly basis.

3. For gas-proportional counters, background measurements shall be performed on at least on a weekly basis.

4. For scintillation counters, background measurements shall be performed each day of use.

E. Instrument contamination monitoring. The laboratory shall have a written procedure for monitoring radiation measurement instrumentation for radioactive contamination. The procedure shall indicate the frequency of the monitoring and shall indicate criteria, which initiates corrective action.

1VAC30-45-805. Radiochemical testing: Minimum detectable activity (MDA)/Minimum detectable concentration (MDC)/Lower level of detection (LLD).

A. MDA/MDC/LLD shall be determined prior to sample analysis and shall be redetermined each time there is a significant change in the test method or instrument type.

B. The procedures employed shall be documented and consistent with mandated method or regulation.

1VAC30-45-806. Radiochemical testing: data reduction.

A. The requirements of 1VAC30-45-730 K apply.

B. Measurement uncertainties. Each result shall be reported with the associated measurement uncertainty. The procedures for determining the measurement uncertainty shall be documented and be consistent with mandated method and regulation.

1VAC30-45-807. Radiochemical testing: quality of standards and reagents.

A. The quality control program shall establish and maintain provisions for radionuclide standards.

1. Reference standards that are used in a radiochemical laboratory shall be obtained from the National Institute of Standards and Technology (NIST), or suppliers who participate in supplying NIST standards or NIST traceable radionuclides. Any reference standards purchased outside the United States shall be traceable back to each country's national standards laboratory. Commercial suppliers of reference standards shall conform to ANSI N42.22 to assure the quality of their products.

2. Reference standards shall be accompanied with a certificate of calibration whose content is as described in ANSI N42.22 - 1995, Section 8, Certificates.

3. Laboratories should consult with the supplier if the laboratory's verification of the activity of the reference traceable standard indicates a noticeable deviation from the certified value. The laboratory shall not use a value other than the decay corrected certified value. The laboratory shall have a written procedure for handling, storing and establishment of expiration dates for reference standards.

B. All reagents used shall be analytical reagent grade or better.

1VAC30-45-808. Radiochemical testing: constant and consistent test conditions.

The laboratory shall maintain a radiological control program that addresses analytical radiological control. The program shall address the procedures for segregating samples with potentially widely varying levels of radioactivity. The radiological control program shall explicitly define how low level and high level samples will be identified, segregated and processed in order to prevent sample cross-contamination. The radiological control program shall include the measures taken to monitor and evaluate background activity or contamination on an ongoing basis.

1VAC30-45-809. (Reserved).

1VAC30-45-810. Air testing: general.

These standards shall apply to samples that are submitted to a laboratory for the purpose of analysis. They do not apply to field activities such as source air emission measurements or the use of continuous analysis devices.

1VAC30-45-811. Air testing: negative and positive controls.

A. Negative controls.

1. Method blanks shall be performed at a frequency of at least one per batch of 20 environmental samples or less per sample preparation method. The results of the method blank analysis shall be used to evaluate the contribution of the laboratory provided sampling media and analytical sample preparation procedures to the amount of analyte found in each sample. If the method blank result is greater than the limit of quantitation and contributes greater than 10% of the total amount of analyte found in the sample, the source of the contamination shall be investigated and measures taken to eliminate the source of contamination. If contamination is found, the data shall be qualified in the report.

2. Collection efficiency. Sampling trains consisting of multiple sections (e.g., filters, sorbent tubes, impingers) that are received intact by the laboratory shall be separated into "front" and "back" sections if required by the client. Each section shall be processed and analyzed separately and the analytical results reported separately.

B. Positive controls. Laboratory control sample (LCS) shall be analyzed at a rate of at least one per batch of 20 or fewer samples per sample preparation method for each analyte. If a spiking solution is not available, a calibration solution whose concentration approximates that of the samples shall be included in each batch and with each lot of media. If a calibration solution must be used for the LCS, the client will be notified prior to the start of analysis. The concentration of the LCS shall be relevant to the intended use of the data and either at a regulatory limit or below it.

C. Surrogates shall be used as required by the test method.

D. Matrix spike shall be used as required by the test method.

1VAC30-45-812. Air testing: analytical variability/reproducibility.

Matrix spike duplicates (MSDs) or laboratory duplicates shall be analyzed at a minimum of one in 20 samples per sample batch. The laboratory shall document their procedure to select the use of appropriate types of spikes and duplicates. The selected samples(s) shall be rotated among sampling points or sampling locations so that various sample matrix problems may be noted and/or addressed. Poor performance in the spikes and duplicates may indicate a problem with the sample composition and shall be reported to the client.

1VAC30-45-813. Air testing: method evaluation.

In order to ensure the accuracy of the reported result, the following procedures shall be in place:

1. Demonstration of capability shall be performed prior to the analysis of any samples and with a significant change in instrument type, personnel, quality system matrix, or test method.

2. Calibration. Calibration protocols specified in 1VAC30-45-740 shall be followed.

3. Proficiency test samples. The results of such analyses shall be used by the laboratory to evaluate the ability of the laboratory to produce accurate data.

1VAC30-45-814. Air testing: limit of detection.

The requirements of 1VAC30-45-771 shall apply.

1VAC30-45-815. Air testing: data reduction.

The procedures for data reduction, such as use of linear regression, shall be documented.

1VAC30-45-816. Air testing: quality of standards and reagents.

A. The source of standards shall comply with 1VAC30-45-740 C.

B. The purity of each analyte standard and each reagent shall be documented by the laboratory through certificates of analyses from the manufacturer/vendor, manufacturer/vendor specifications, and/or independent analysis.

C. In methods where the purity of reagents is not specified, analytical reagent grade or higher quality, if available, shall be used.

1VAC30-45-817. Air testing: selectivity.

The laboratory shall develop and document acceptance criteria for test method selectivity such as absolute and relative retention times, wavelength assignments, mass spectral library quality of match, and mass spectral tuning.

1VAC30-45-818. Air testing: constant and consistent test conditions.

A. The laboratory shall assure that the test instruments consistently operate within the specifications required of the application for which the equipment is used.

B. The laboratory shall document that all sampling equipment, containers and media used or supplied by the laboratory meet required test method criteria.

C. If supplied or used by the laboratory, procedures for field equipment decontamination shall be developed and their use documented.

D. The laboratory shall have a documented program for the calibration and verification of sampling equipment such as pumps, meter boxes, critical orifices, flow measurement devices and continuous analyzers, if these equipment are used or supplied by the laboratory.

1VAC30-45-819. (Reserved).

1VAC30-45-820. Asbestos testing: general.

These standards apply to laboratories undertaking the examination of asbestos samples. These standards are organized by analytical technique, including transmission electron microscopy (TEM) for the analysis of water, wastewater, air, and bulk samples; phase contrast microscopy (PCM) for analysis of workplace air; and polarized light microscopy (PLM) for analysis of bulk samples. These procedures for asbestos analysis involve sample preparation followed by detection of asbestos. If NIST SRMs specified below are unavailable, the laboratory may substitute an equivalent reference material with a certificate of analysis.

1VAC30-45-821. Asbestos testing: negative controls.

A. Transmission electron microscopy.

1. Water and wastewater.

a. Blank determinations shall be made prior to sample collection. When using polyethylene bottles, one bottle from each batch, or a minimum of one from each 24 shall be tested for background level. When using glass bottles, four bottles from each 24 shall be tested. An acceptable bottle blank level is defined as ≤0.01 MFL > 10µm. (EPA/600/R-94/134, Method 100.2, Section 8.2)

b. A process blank sample consisting of fiber-free water shall be run before the first field sample. The quantity of water shall be ≥10 mL for a 25-mm diameter filter and ≥ 50 mL for a 47-mm diameter filter. (EPA /600/R-94/134, Method 100.2, Section 11.8)

2. Air.

a. A blank filter shall be prepared with each set of samples. A blank filter shall be left uncovered during preparation of the sample set and a wedge from that blank filter shall be prepared alongside wedges from the sample filters. At minimum, the blank filter shall be analyzed for each 20 samples analyzed. (40 CFR Part 763, Appendix A to Subpart E (AHERA), Table 1)

b. Maximum contamination on a single blank filter shall be no more than 53 structures/mm2. Maximum average contamination for all blank filters shall be no more than 18 structures/mm2. (AHERA, III.F.2)

3. Bulk samples.

a. Contamination checks using asbestos-free material, such as the glass fiber blank in SRM 1866 (Page C-3, NIST Handbook 150-3, August 1994) shall be performed at a frequency of one for every 20 samples analyzed. The detection of asbestos at a concentration exceeding 0.1% will require an investigation to detect and remove the source of the asbestos contamination.

b. The laboratory shall maintain a list of nonasbestos fibers that can be confused with asbestos (Section 7.5, Page C-8, NIST Handbook 150-3, August 1994). The list shall include crystallographic and/or chemical properties that disqualify each fiber being identified as asbestos (Section 2.5.5.2.1 Identification, Page 54, EPA/600/R-93/116).

c. The laboratory should have a set of reference asbestos materials from which a set of reference diffraction and X-ray spectra have been developed.

B. Phase contrast microscopy.  At least two field blanks (or 10% of the total samples, whichever is greater) shall be submitted for analysis with each set of samples. Field blanks shall be handled in a manner representative of actual handling of associated samples in the set with a single exception that air shall not be drawn through the blank sample. A blank cassette shall be opened for approximately 30 seconds at the same time other cassettes are opened just prior to analysis. Results from field blank samples shall be used in the calculation to determine final airborne fiber concentration. The identity of blank filters should be unknown to the counter until all counts have been completed. If a field blank yields greater than seven fibers per 100 graticule fields, report possible contamination of the samples.

C. Polarized light microscopy.

1. Friable materials. At least one blank slide shall be prepared daily or with every 50 samples analyzed, whichever is less. This is prepared by mounting a subsample of an isotropic verified non-ACM (e.g., fiberglass in SRM 1866) in a drop of immersion oil (nD should reflect usage of various nD’s) on a clean slide, rubbing preparation tools (forceps, dissecting needles, etc.) in the mount and placing a clean coverslip on the drop. The entire area under the coverslip shall be scanned to detect any asbestos contamination. A similar check shall be made after every 20 uses of each piece of homogenization equipment. An isotropic verified non-ACM shall be homogenized in the clean equipment, a slide prepared with the material and the slide scanned for asbestos contamination. (This can be substituted for the blank slide mentioned in this section.)

2. Nonfriable materials. At least one non-ACM nonfriable material shall be prepared and analyzed with every 20 samples analyzed. This non-ACM shall go through the full preparation and analysis regimen for the type of analysis being performed.

1VAC30-45-822. Asbestos testing: test variability/reproducibility.

A. Transmission electron microscopy. Quality assurance analyses shall be performed regularly covering all time periods, instruments, tasks, and personnel. The selection of samples shall be random and samples of special interest may be included in the selection of samples for quality assurance analyses. When possible, the checks on personnel performance shall be executed without their prior knowledge. A disproportionate number of analyses shall not be performed prior to internal or external audits. It is recommended that a laboratory initially be at 100% quality control (all samples reanalyzed). The proportion of quality control samples can later be lowered gradually, as control indicates, to a minimum of 10%.

1. Water and wastewater. All analyses shall be performed on relocator grids so that other laboratories can easily repeat analyses on the same grid openings. Quality assurance analyses shall not be postponed during periods of heavy workloads. The total number of QA samples and blanks shall be greater than or equal to 10% of the total sample workload. Precision of analyses is related to concentration, as gleaned from interlaboratory proficiency testing. Relative standard deviations (RSD) for amphibole asbestos decreased from 50% at 0.8 MFL to 25% at 7 MFL in interlaboratory proficiency testing, while RSD for chrysotile was higher, 50% at 6 MFL.

a. Replicate. A second, independent analysis shall be performed on the same grids but on different grid openings than used in the original analysis of a sample. Results shall be within 1.5X of Poisson standard deviation. This shall be performed at a frequency of 1 per 100 samples. (EPA/600/R-94/134, Method 100.2, Table 2)

b. Duplicate. A second aliquot of sample shall be filtered through a second filter, prepared and analyzed in the same manner as the original preparation of that sample. Results shall be within 2.0X of Poisson standard deviation. This shall be performed at a frequency of one per 100 samples. (EPA/600/R-94/134, Method 100.2, Table 2)

c. Verified analyses. A second, independent analysis shall be performed on the same grids and grid openings used in the original analysis of a sample. The two sets of results shall be compared according to Turner and Steel (NISTIR 5351). This shall be performed at a frequency of one per 20 samples. Qualified analysts shall maintain an average of ≥ 80% true positives, ≤ 20% false negatives, and ≤ 10% false positives.

2. Air.

a. All analyses shall be performed on relocator grids so that other laboratories can easily repeat analyses on the same grid openings.

b. The laboratory and TEM analysts shall obtain mean analytical results on NIST SRM 1876b so that trimmed mean values fall within 80% of the lower limit and 110% of the upper limit of the 95% confidence limits as published on the certificate. These limits are derived from the allowable false positives and false negatives given in subdivision A 2 e (3) of this subsection. SRM 1876b shall be analyzed a minimum of once per year by each TEM analyst.

c. The laboratory shall have documentation demonstrating that TEM analysts correctly classify at least 90% of both bundles and single fibrils of asbestos structures greater than or equal to 1 mm in length in known standard materials traceable to NIST, such as NIST bulk asbestos SRM 1866.

d. Interlaboratory analyses shall be performed to detect laboratory bias. The frequency of interlaboratory verified analysis shall correspond to a minimum of 1 per 200 grid square analyses.

e. If more than one TEM is used for asbestos analysis, intermicroscope analyses shall be performed to detect instrument bias.

(1) Replicate. A second, independent analysis shall be performed in accordance with Section D.6.2.1.1.a. (AHERA, Table III)

(2) Duplicate. A second wedge from a sample filter shall be prepared and analyzed in the same manner as the original preparation of that sample. Results shall be within 2.0X of Poisson standard deviation. This shall be performed at a frequency of 1 per 100 samples. (AHERA, Table III)

(3) Verified analyses. A second, independent analysis shall be performed on the same grids and grid openings in accordance with subdivision A 1 c of this section.

3. Bulk samples. Determination of precision and accuracy should follow guidelines in NISTIR 5951, Guide for Quality Control on the Qualitative and Quantitative Analysis of Bulk Asbestos Samples: Version 1. Because bulk samples with low (< 10%) asbestos content are the most problematic, a laboratory’s quality control program should focus on such samples. At least 30% of a laboratory’s QC analyses shall be performed on samples containing from 1.0% to 10% asbestos.

(a) Intra-analyst precision. At least one out of 50 samples shall be reanalyzed by the same analyst. For single analyst laboratories, at least one out of every 10 samples shall be reanalyzed by the same analyst.

(b) Inter-analyst precision. At least one out of 15 samples shall be reanalyzed by another analyst. Inter-analyst results will require additional reanalysis, possibly including another analyst, to resolve discrepancies when classification (ACM vs. non-ACM) errors occur, when asbestos identification errors occur, or when inter-analyst precision is found to be unacceptable.

(c) Inter-laboratory precision. The laboratory shall participate in round robin testing with at least one other laboratory. Samples shall be sent to this other lab at least four times per year. These samples shall be samples previously analyzed as QC samples. Results of these analyses shall be assessed in accordance with QC requirements. As a minimum, the QC requirements shall address misclassifications (false positives, false negatives) and misidentification of asbestos types.

B. Phase contrast microscopy.

1. Inter-laboratory precision. Each laboratory analyzing air samples for compliance determination shall implement an inter-laboratory quality assurance program that as a minimum includes participation of at least two other independent laboratories. Each laboratory shall participate in round robin testing at least once every six months with at least all the other laboratories in its inter-laboratory quality assurance group. Each laboratory shall submit slides typical of its own workload for use in this program. The round robin shall be designed and results analyzed using appropriate statistical methodology. Results of this QA program shall be posted in each laboratory to keep the microscopists informed.

2. Intra- and inter-analyst precision. Each analyst shall select and count a prepared slide from a "reference slide library" on each day on which air counts are performed. Reference slides shall be prepared using well-behaved samples taken from the laboratory workload. Fiber densities shall cover the entire range routinely analyzed by the laboratory. These slides shall be counted by all analysts to establish an original standard deviation and corresponding limits of acceptability. Results from the daily reference sample analysis shall be compared to the statistically derived acceptance limits using a control chart or a database. It is recommended that the labels on the reference slides be periodically changed so that the analysts do not become familiar with the samples. Intra- and inter-analyst precision may be estimated from blind recounts on reference samples. Inter-analyst precision shall be posted in each laboratory to keep the microscopists informed.

C. Polarized light microscopy. Refer to subdivision A 3 of this section.

1VAC30-45-823. Asbestos testing: other quality control measures.

A. Transmission electron microscopy.

1. Water and wastewater.

a. Filter preparations shall be made from all six asbestos types from NIST SRMs 1866 and 1867. These preparations shall have concentrations between one and 20 structures (> 10µm) per 0.01 mm2. One of these preparations shall be analyzed independently at a frequency of one per 100 samples analyzed. Results shall be evaluated as verified asbestos analysis in accordance with Turner and Steel (NISTIR 5351).

b. NIST SRM 1876b shall be analyzed annually by each analyst. Results shall be evaluated in accordance with limits published for that SRM. This SRM is not strictly appropriate for waterborne asbestos but analysts can demonstrate general TEM asbestos competence by producing results within the published limits of this (the only recognized TEM counting standard) SRM.

2. Air

a. Filter preparations shall be made from all six asbestos types in accordance with subdivision A 1 a of this section.

b. NIST SRM 1876b shall be analyzed annually in accordance with subdivision A 1 b of this section.

3. Bulk samples. All analysts shall be able to correctly identify the six regulated asbestos types (chrysotile, amosite, crocidolite, anthophyllite, actinolite, and tremolite). Standards for the six asbestos types listed are available from NIST (SRMs 1866 and 1867). These materials can also be used as identification standards for AEM (Section 3.2.1 Qualitative Analysis, Page 57, EPA/600/R-93/116).

B. Phase contrast microscopy.

1. Test for nonrandom fiber distribution. Blind recounts by the same analyst shall be performed on 10% of the filters counted. A person other than the counter should re-label slides before the second count. A test for type II error (NIOSH 7400, Issue 2, 15 August 1994, Section 13) shall be performed to determine whether a pair of counts by the same analyst on the same slide should be rejected due to nonrandom fiber distribution. If a pair of counts is rejected by this test, the remaining samples in the set shall be recounted and the new counts shall be tested against first counts. All rejected paired counts shall be discarded. It shall not be necessary to use this statistic on blank recounts.

2. All individuals performing airborne fiber analysis shall have taken the NIOSH Fiber Counting Course for sampling and evaluating airborne asbestos dust or an equivalent course.

3. All laboratories shall participate in a national sample testing scheme such as the Proficiency Analytical Testing (PAT) program or the Asbestos Analysts Registry (AAR) program, both sponsored by the American Industrial Hygiene Association (AIHA), or equivalent.

C. Polarized light microscopy.

1. Friable materials. Because accuracy cannot be determined by reanalysis of routine field samples, at least one out of 100 samples shall be a standard or reference sample that has been routinely resubmitted to determine analyst’s precision and accuracy. A set of these samples should be accumulated from proficiency testing samples with predetermined weight compositions or from standards generated with weighed quantities of asbestos and other bulk materials (Perkins and Harvey, 1993; Parekh et al., 1992; Webber et al., 1982). At least half of the reference samples submitted for this QC shall contain between 1.0% and 10% asbestos.

2. Nonfriable materials. At least one out of 100 samples shall be a verified quantitative standard that has routinely been resubmitted to determine analyst precision and accuracy.

1VAC30-45-824. Asbestos testing: method evaluation.

In order to ensure the accuracy of reported results, the following procedures shall be in place:

1. Demonstration of capability shall be performed initially (prior to the analysis of any samples) and with a significant change in instrument type, personnel, or method.

2. Performance audits. The results of such analyses shall be used by the laboratory to evaluate the ability of the laboratory to produce accurate data.

1VAC30-45-825. Asbestos testing: asbestos calibration.

Refer to methods referenced in the following sections for specific equipment requirements.

1. Transmission electron microscopy: general. Analytical electron microscopy equipment will not be discussed in this document.

2. Transmission electron microscopy: water and wastewater. All calibrations listed below (unless otherwise noted) shall be performed under the same analytical conditions used for routine asbestos analysis and shall be recorded in a notebook and include date and analyst’s signature. Frequencies stated below may be reduced to "before next use" if no samples are analyzed after the last calibration period has expired. Likewise, frequencies may have to be increased following non-routine maintenance or unacceptable calibration performance.

a. Magnification calibration. Magnification calibration shall be done at the fluorescent screen, with the calibration specimen at the eucentric position, at the magnification used for fiber counting, generally 10,000 and 20,000x. A logbook shall be maintained with the dates of the calibration recorded. Calibrations shall be performed monthly to establish the stability of magnification. Calibration data shall be displayed on control charts that show trends over time. (EPA/600/R-94/134, Method 100.2, Section 10.1)

b. Camera constant. The camera length of the TEM in the Selected Area Electron Diffraction (SAED) mode shall be calibrated before SAED patterns of unknown samples are observed. The diffraction specimen shall be at the eucentric position for this calibration. This calibration shall allow accurate (< 10% variation) measurement of layer-line spacings on the medium used for routine measurement, i.e., the phosphor screen or camera film. This shall also allow accurate (< 5.0% variation) measurement of zone axis SAED patterns on permanent media, e.g., film. Calibrations shall be performed monthly to establish the stability of the camera constant (EPA/600/R-94/134, Method 100.2, Section 10.2). Where nonasbestiform minerals may be expected (e.g., winchite, richterite, industrial talc, vermiculite, etc.), an internal camera constant standard such as gold, shall be deposited and measured on each sample to facilitate accurate indexing of zone axis SAED patterns. In such cases, layer line analysis alone shall not be used. Calibration data shall be displayed on control charts that show trends over time.

c. Spot size. The diameter of the smallest beam spot at crossover shall be less than 250 nm as calibrated quarterly. Calibration data shall be displayed on control charts that show trends over time. (EPA /600/R-94/134, Method 100.2, Section 10.3)

d. Beam dose. The beam dose shall be calibrated so that beam damage to chrysotile is minimized, specifically so that an electron diffraction pattern from a single fibril ≥1 µm in length from a NIST SRM chrysotile sample is stable in the electron beam dose for at least 15 seconds.

e. EDXA system.

(1) The x-ray energy vs. channel number for the EDXA system shall be calibrated to within 20 eV for at least two peaks between 0.7 keV and 10 keV. One peak shall be from the low end (0.7 keV to 2 keV) and the other peak from the high end (7 keV to 10 keV) of this range. The calibration of the x-ray energy shall be checked prior to each analysis of samples and recalibrated if out of the specified range.

(2) The ability of the system to resolve the Na Ka line from the Cu L line shall be confirmed quarterly by obtaining a spectrum from the NIST SRM 1866 crocidolite sample on a copper grid.

(3) The k-factors for elements found in asbestos (Na, Mg, Al, Si, Ca, and Fe) relative to Si shall be calibrated semiannually, or anytime the detector geometry may be altered. NIST SRM 2063a shall be used for Mg, Si, Ca, Fe, while k-factors for Na and Al may be obtained from suitable materials such as albite, kaersutite, or NIST SRM 99a. The k-factors shall be determined to a precision (2s) within 10% relative to the mean value obtained for Mg, Al, Si, Ca, and Fe, and within 20% relative to the mean value obtained for Na. The k-factor relative to Si for Na shall be between 1.0 and 4.0, for Mg and Fe shall be between 1.0 and 2.0, and for Al and Ca shall be between 1.0 and 1.75. The k-factor for Mg relative to Fe shall be 1.5 or less. Calibration data shall be displayed on control charts that show trends over time.

(4) The detector resolution shall be checked quarterly to ensure a full-width half-maximum resolution of <175 eV at Mn Ka (5.90 keV). Calibration data shall be displayed on control charts that show trends over time.

(5) The portions of a grid in a specimen holder for which abnormal x-ray spectra are generated under routine asbestos analysis conditions shall be determined and these areas shall be avoided in asbestos analysis.

(6) The sensitivity of the detector for collecting x-rays from small volumes shall be documented quarterly by collecting resolvable Mg and Si peaks from a unit fibril of NIST SRM 1866 chrysotile.

f. Low temperature asher. The low temperature asher shall be calibrated quarterly by determining a calibration curve for the weight vs. ashing time of collapsed mixed-cellulose-ester (MCE) filters. Calibration data shall be displayed on control charts that show trends over time.

g. Grid openings. The magnification of the grid opening measurement system shall be calibrated using an appropriate standard at a frequency of 20 openings/20 grids/lot of 1000 or one opening/sample. The variation in the calibration measurements (2s) is <5.0% of the mean calibration value.

3. Air. All calibrations shall be performed in accordance with subdivision 2 of this section, with the exception of magnification. Magnification calibration shall be done at the fluorescent screen, with the calibration specimen at the eucentric position, at the magnification used for fiber counting, generally 15,000 to 20,000x (AHERA, III.G.1.c). A logbook shall be maintained with the dates of the calibration recorded. Calibrations shall be performed monthly to establish the stability of magnification.

4. Bulk samples. All calibrations shall be performed in accordance with subdivision 3 of this section.

5. Phase contrast microscopy.

a. At least once daily, the analyst shall use the telescope ocular (or Bertrand lens, for some microscopes) supplied by the manufacturer to ensure that the phase rings (annular diaphragm and phase-shifting elements) are concentric.

b. The phase-shift limit of detection of the microscope shall be checked monthly or after modification or relocation using an HSE/NPL phase-contrast test slide for each analyst/microscope combination (refer to NIOSH 7400, Issue 2, 15 August 1994, Section 10b). This procedure assures that the minimum detectable fiber diameter (< ca. 0.25mm) for this microscope is achieved.

c. Prior to ordering the Walton-Beckett graticule, calibration, in accordance with NIOSH 7400, Issue 2, 15 August 1994, Appendix A, shall be performed to obtain a counting area 100 mm in diameter at the image plane. The diameter, dc (mm), of the circular counting area and the disc diameter shall be specified when ordering the graticule. The field diameter (D) shall be verified (or checked), to a tolerance of 100 µm ± 2 µm, with a stage micrometer upon receipt of the graticule from the manufacturer. When changes (zoom adjustment, disassembly, replacement, etc.) occur in the eyepiece-objective-reticle combination, field diameter shall be remeasured (or recalibrated) to determine field area (mm2). Recalibration of field diameter shall also be required when there is a change in interpupillary distance (i.e., change in analyst). Acceptable range for field area shall be 0.00754 mm2 to 0.00817 mm2. The actual field area shall be documented and used.

6. Polarized light microscopy.

a. Microscope alignment. To accurately measure the required optical properties, a properly aligned polarized light microscope (PLM) shall be utilized. The PLM shall be aligned before each use. (Section 2.2.5.2.3, EPA/600/R-93/116, July 1993)

b. Refractive index liquids. Series of nD = 1.49 through 1.72 in intervals less than or equal to 0.005. Refractive index liquids for dispersion staining, high-dispersion series 1.550, 1.605, 1.680. The accurate measurement of the refractive index (RI) of a substance requires the use of calibrated refractive index liquids. These liquids shall be calibrated at first use and semiannually, or next use, whichever is less frequent, to an accuracy of 0.004, with a temperature accuracy of 2°C using a refractometer or RI glass beads.

1VAC30-45-826. Asbestos testing: analytical sensitivity.

A. Transmission electron microscopy.

1. Water and wastewater. An analytical sensitivity of 200,000 fibers per liter (0.2 MFL) is required for each sample analyzed (EPA/600/R-94/134, Method 100.2, Section 1.6). Analytical sensitivity is defined as the waterborne concentration represented by the finding of one asbestos structure in the total area of filter examined. This value will depend on the fraction of the filter sampled and the dilution factor (if applicable).

2. Air. An analytical sensitivity of 0.005 structures/cm2 is required for each sample analyzed. Analytical sensitivity is defined as the airborne concentration represented by the finding of one asbestos structure in the total area of filter examined. This value will depend on the effective surface area of the filter, the filter area analyzed, and the volume of air sampled (AHERA, Table I).

3. Bulk samples.

a. The range is dependent on the type of bulk material being analyzed. The sensitivity may be as low as 0.0001% depending on the extent to which interfering materials can be removed during the preparation of AEM specimens. (Section 2.5.2 Range, Page 51, EPA/600/R-93/116)

b. There should be an error rate of less than 1.0% on the qualitative analysis for samples that contain chrysotile, amosite, and crocidolite. A slightly higher error rate may occur for samples that contain anthophyllite, actinolite, and tremolite, as it can be difficult to distinguish among the three types. (Section 3, Page 10, NIST Handbook 150-3, August 1994)

B. Phase contrast microscopy. The normal quantitative working range of the test method is 0.04 to 0.5 fiber/ cm2 for a 1000 L air sample. An ideal counting range on the filter shall be 100 to 1300 fibers/mm2. The limit of detection (LOD) is estimated to be 5.5 fibers per 100 fields or 7 fibers/mm2. The LOD in fiber/cc will depend on sample volume and quantity of interfering dust but shall be <0.01 fiber/ cm2 for atmospheres free of interferences. (NIOSH 7400, Issue 2, 15 August 1994)

C. Polarized light microscopy. The laboratory shall utilize a test method that provides a limit of detection that is appropriate and relevant for the intended use of the data. Limit of detection shall be determined by the protocol in the test method or applicable regulation.

1VAC30-45-827. Asbestos testing: data reduction.

A. Transmission electron microscopy.

1. Water and wastewater.

a. The concentration of asbestos in a given sample shall be calculated in accordance with EPA/600/R-94/134, Method 100.2, Section 12.1. Refer to 1VAC30-45-730 K for additional data reduction requirements.

b. Measurement uncertainties. The laboratory shall calculate and report the upper and lower 95% confidence limits on the mean concentration of asbestos fibers found in the sample (EPA /600/R-94/134, Method 100.2, Section 12.2.2).

2. Air.

a. The concentration of asbestos in a given sample shall be calculated in accordance with the method utilized, e.g., AHERA. Refer to 1VAC30-45-730 K for additional data reduction requirements.

b. Measurement uncertainties. The laboratory shall calculate and report the upper and lower 95% confidence limits on the mean concentration of asbestos fibers found in the sample.

3. Bulk samples.

a. The concentration of asbestos in a given sample shall be calculated in accordance with the method utilized (e.g., EPA/600/R-93/116, July 1993). Refer to 1VAC30-45-730 K for additional data reduction requirements.

b. Measurement uncertainties. Proficiency testing for floor tiles analyzed by TEM following careful gravimetric reduction (New York ELAP Certification Manual Item 198.4) has revealed an interlaboratory standard deviation of approximately 20% for residues containing 70% or more asbestos. Standard deviations range from 20% to 60% for residues with lower asbestos content.

B. Phase contrast microscopy.

1. Airborne fiber concentration in a given sample shall be calculated in accordance with NIOSH 7400, Issue 2, 15 August 1994, Sections 20 and 21. Refer to 1VAC30-45-730 K for additional data reduction requirements.

2. Measurement uncertainties. The laboratory shall calculate and report the intra-laboratory and inter-laboratory relative standard deviation with each set of results. (NIOSH 7400, Issue 2, 15 August 1994)

3. Fiber counts above 1300 fibers/mm2 and fiber counts from samples with >50% of the filter area covered with particulate should be reported as "uncountable" or "probably biased." Other fiber counts outside the 100-1300 fibers/mm2 range should be reported as having "greater than optimal variability" and as being "probably biased."

C. Polarized light microscopy.

1. The concentration of asbestos in a given sample shall be calculated in accordance with the method utilized (e.g., EPA/600/R-93/116, July 1993). Refer to 1VAC30-45-730 K for additional data reduction requirements.

2. Method uncertainties. The individual laboratory shall determine precision and accuracy for the percent range involved. If point counting and/or visual estimates are used, a table of reasonable expanded errors (refer to EPA/600/R-93/116, July 1993, Table 2-1) should be generated for different concentrations of asbestos.

1VAC30-45-828. Asbestos testing: quality of standards and reagents.

A. Transmission electron microscopy.

1. The quality control program shall establish and maintain provisions for asbestos standards.

a. Reference standards that are used in an asbestos laboratory shall be obtained from the National Institute of Standards and Technology (NIST), EPA, or suppliers who participate in supplying NIST standards or NIST traceable asbestos. Any reference standards purchased outside the United States shall be traceable back to each country’s national standards laboratory. Commercial suppliers of reference standards shall conform to ANSI N42.22 to assure the quality of their products.

b. Reference standards shall be accompanied with a certificate of calibration whose content is as described in ANSI N42.22-1995, Section 8, Certificates.

2. All reagents used shall be analytical reagent grade or better.

3. The laboratory shall have mineral fibers or data from mineral fibers that will allow differentiating asbestos from at least the following "look-alikes": fibrous talc, sepiolite, wollastonite, attapulgite (palygorskite), halloysite, vermiculite scrolls, antigorite, lizardite, pyroxenes, hornblende, richterite, winchite, or any other asbestiform minerals that are suspected as being present in the sample.

B. Phase contrast microscopy. Standards of known concentration have not been developed for this testing method. Routine workload samples that have been statistically validated and national proficiency testing samples such as PAT and AAR samples available from the AIHA may be utilized as reference samples (refer to 1VAC30-45-822 B 2) to standardize the optical system and analyst. All other testing reagents and devices (HSE/NPL test slide and Walton-Beckett Graticule) shall conform to the specifications of the method (refer to NIOSH 7400, Issue 2, 15 August 1994).

C. Polarized light microscopy. Refer to 1VAC30-45-828 A.

1VAC30-45-829. Asbestos testing: constant and consistent test conditions.

The laboratory shall establish and adhere to written procedures to minimize the possibility of cross-contamination between samples.

1VAC30-45-830. (Reserved).

1VAC30-45-840. (Reserved).

[ 1VAC30-45-760 1VAC30-45-850 ]. Sample handling, sample acceptance policy and sample receipt.

While the laboratory may not have control of field sampling activities, the following are essential to ensure the validity of the laboratory’s data.

1. Sample tracking. The laboratory shall have a documented system for uniquely identifying the items to be tested to ensure that there can be no confusion regarding the identity of such items at any time.  This system shall include identification for all samples, subsamples and subsequent extracts or digestates or both.  The use of container shape, size or other physical characteristic, such as amber glass or purple top, is not an acceptable means of identifying the sample. [ System laboratories shall use a permanent chronological record such as a logbook or electronic database to document receipt of all containers. This sample receipt log shall record the following at a minimum: name of facility where sample was taken, date and time of laboratory receipt, unique laboratory ID code, and signature or initials of the person making the entries. ]

2. Sample acceptance policy. The laboratory shall have a written sample acceptance policy that clearly outlines the circumstances under which samples shall be accepted or rejected. The policy shall ensure that only properly obtained samples [ with appropriate sampling records (see 1VAC30-45-640 B) ] are analyzed and that the samples are handled properly.  This sample acceptance policy shall be made available to sample collection personnel.  The policy shall include elements such as appropriate documentation of the sample’s identification, use of appropriate sample containers, adherence to specified holding times, adequate sample volume to perform necessary tests, and procedures to be used when samples show signs of damage, contamination or inadequate preservation.

3. Sample receipt protocols.

a. Upon receipt, the condition of the sample, including any abnormalities or departures from standard condition as prescribed in the relevant test method, shall be recorded.  All items specified by the sample acceptance policy shall be checked.

b. All samples that require thermal preservation shall be considered acceptable if the arrival temperature is either within 2 degrees Celsius of the required temperature or the method specified range.  For samples with a specified temperature of 4 degrees Celsius, samples with a temperature of ranging from just above freezing temperature of water to 6 degrees Celsius shall be acceptable.  Samples that are hand delivered to the laboratory immediately after collection [ or on the same day that are collected ] may not meet this criteria.  In these cases, the samples shall be considered acceptable if there is evidence that the chilling process has begun such as arrival on ice.

c. The laboratory shall implement procedures for checking chemical preservation using readily available techniques, such as pH or free chlorine prior to or during sample preparation or analysis.

d. The results of all checks required by the sample acceptance policy and relevant test method shall be recorded.

4. Storage conditions.

a. The laboratory shall have documented procedures and appropriate facilities to avoid deterioration, contamination or damage to the sample during storage, handling, preparation, and testing.  Any relevant instructions provided with the item shall be followed.  Where items have to be stored or conditioned under specific environmental conditions, these conditions shall be maintained, monitored and recorded.

b. Samples shall be stored according to the conditions specified by preservation protocols:

(1) Samples that require thermal preservation shall be stored under refrigeration that is within 2 degrees Celsius of the specified preservation temperature unless method specific criteria exist.  For samples with a specified storage temperature of 4 degrees Celsius, storage at a temperature above the freezing point of water to 6 degrees Celsius shall be acceptable.

(2) Samples shall be stored away from all standards, reagents, food and other potentially contaminating sources.  Samples shall be stored in such a manner to prevent cross contamination.

c. Sample fractions, extracts, leachates and other sample preparation products shall be stored according to subdivision 4 a of this section or according to specifications in the test method.

d. Where a sample or portion of the sample is to be held secure (for example, for reasons of record, safety or value, or to enable check calibrations or tests to be performed later), the laboratory shall have storage and security arrangements that protect the condition and integrity of the secured items or portions concerned.

5. Sample disposal.  The laboratory shall have standard operating procedures for the disposal of samples, digestates, leachates and extracts or other sample preparation products.

[ 1VAC30-45-770 1VAC30-45-860 ]. Laboratory report format and contents.

A. The results of each test or series of tests carried out by the laboratory shall be reported accurately, clearly, unambiguously and objectively. The results shall normally be reported in a test report required by regulation and shall include all the information necessary for the interpretation of the test results and all information required by the method used.

B. Where the certificate or report contains results of tests performed by subcontractors, these results shall be clearly identified by subcontractor name or applicable certification number.

C. After issuance of the report, the laboratory report shall remain unchanged. Material amendments to a calibration certificate, test report or test certificate after issue shall be made only in the form of a further document, or data transfer including the statement "Supplement to Test Report or Test Certificate, serial number . . . (or as otherwise identified)," or equivalent form of wording. Such amendments shall meet all the relevant requirements of this article.

NOTICE: The forms used in administering the above regulation are not being published; however, the name of each form is listed below. The forms are available for public inspection by contacting the agency contact for this regulation, or at the office of the Registrar of Regulations, General Assembly Building, 2nd Floor, Richmond, Virginia.

FORMS

Application for Certification of Environmental Laboratories [ Performing Only Simple Test Procedures Under 1VAC35-45 (with instructions), eff. xx/xx DGS/DCLS 1 (eff. 10/08) ].

[ Application for Certification of General Environmental Laboratories under 1VAC35-45 (with instructions), eff. xx/xx. ]

CHAPTER 46 
ACCREDITATION FOR COMMERCIAL ENVIRONMENTAL LABORATORIES

Part I 
General Provisions

1VAC30-46-10. Purpose.

Section 2.2-1105 of the Code of Virginia directs the Division of Consolidated Laboratory Services to establish a program to certify environmental laboratories that perform tests, analyses, measurements or monitoring required  [ by pursuant to ] the Commonwealth's air, waste and water laws and regulations. This chapter sets out the required standards and the process by which owners [ or operators ] of commercial environmental laboratories may obtain certification for their laboratories. Certification is referred to as accreditation in this chapter. Commercial environmental laboratories are accredited under the standards of the National Environmental Laboratory Accreditation Conference [ (NELAC) ] as approved in [ 2002 2003 ]. In addition, this chapter sets out the process that [ NELAC NELAP ] -accredited environmental laboratories [ located outside Virginia ] must use to receive accreditation in Virginia. 1VAC30-45 covers noncommercial environmental laboratories.

1VAC30-46-20. Establishment of accreditation program.

A. Once the accreditation program has been established, laboratory accreditation shall be required before any environmental analyses performed by a commercial environmental laboratory may be used for the purposes of the Virginia Air Pollution Control Law, the Virginia Waste Management Act or the State Water Control Law (§10.1-1300 et seq., §10.1-1400 et seq., and §62.1-44.2 et seq., respectively, of the Code of Virginia).

B. The accreditation program shall be established on [ the first day of the 25th month following the effective date of this chapter October 1, 2011 ].

1VAC30-46-30. Applicability.

A. General applicability. This chapter applies to the following:

1. Any owner [ or operator ] of a commercial environmental laboratory.

2. Any owner [ or operator ] of an environmental laboratory [ located in jurisdictions outside of Virginia holding NELAP accreditation from a primary accrediting authority ] who wishes to apply for reciprocal accreditation under 1VAC30-46-140.

B. DGS-DCLS.

1. NELAP-accredited laboratory. DGS-DCLS shall meet the requirements of this chapter through review and accreditation by a NELAP-accredited federal or state accrediting authority. This process shall be completed before the program under this chapter and 1VAC30-45 is established.

2. Primary accrediting authority. DGS-DCLS shall meet the requirements of the NELAC [ Standards standards ] to become the primary accrediting authority for the Commonwealth of Virginia. This review and approval by a NELAP accrediting team shall be completed no later than [ one year following the effective date of this chapter October 1, 2009 ].

C. Voluntary accreditation. Any owner [ or operator ] of an environmental laboratory may apply for accreditation under this chapter.

D. Environmental laboratories required to obtain drinking water certification under 1VAC30-40. Any owner [ or operator ] of an environmental laboratory who must meet the requirements of 1VAC30-40 pertaining to drinking water laboratory certification and either 1VAC30-45 or this chapter may meet those requirements by obtaining accreditation under this chapter.

1VAC30-46-40. Definitions.

[ The definitions in the 2003 National Environmental Laboratory Accreditation Conference (NELAC) standards, Chapter 1, Appendix A – Glossary, are incorporated by reference into this section. Some of the definitions from this glossary are included in this section because the terms are used throughout this chapter. Where a term is defined in this section, the term shall have no other meaning, even if it is defined differently in the Code of Virginia or another regulation of the Virginia Administrative Code. Unless specifically defined in this section, the terms used in this chapter shall have the meanings commonly ascribed to them by recognized authorities. ]

"Accreditation" means [ the process by which an agency or organization evaluates and recognizes a laboratory as meeting certain predetermined qualifications or standards, thereby accrediting the laboratory. "Accreditation" is ] the term used as a substitute for the term "certification" under this chapter.

"Accrediting authority" means the territorial, state, or federal agency having responsibility and accountability for environmental laboratory accreditation and which grants accreditation [ under NELAC ].

"Acceptance criteria" means specified limits placed on characteristics of an item, process, or service defined in requirement documents.

[ "Algae" means simple single-celled, colonial, or multicelled, mostly aquatic plants, containing chlorophyll and lacking roots, stems and leaves that are either suspended in water (phytoplankton) or attached to rocks and other substrates (periphyton). ]

"Analyte" means the substance or physical property to be determined in samples examined.

"Analytical method" means a technical procedure for providing analysis of a sample, defined by a body such as the Environmental Protection Agency or the American Society for Testing and Materials, that may not include the sample preparation method.

"Assessment" means the evaluation process used to measure or establish the performance, effectiveness, and conformance of an organization and its systems or both to defined criteria.

"Assessor" means the person who performs on-site assessments of laboratories' capability and capacity for meeting the requirements under this chapter by examining the records and other physical evidence for each one of the tests for which accreditation has been requested.

"Authority" means, in the context of a governmental body or local government, an authority created under the provisions of the Virginia Water and Waste Authorities Act, Chapter 51 (§15.2-5100 et seq.) of Title 15.2 of the Code of Virginia.

[ "Benthic macroinvertebrates" means bottom dwelling animals without backbones that live at least part of their life cycles within or upon available substrates within a body of water. ]

"Commercial environmental laboratory" means an environmental laboratory where environmental analysis is performed for another person.

"Corrective action" means the action taken to eliminate the causes of an existing nonconformity, defect or other undesirable situation in order to prevent recurrence.

"DGS-DCLS" means the Division of Consolidated Laboratory Services of the Department of General Services.

"Environmental analysis" or "environmental analyses" means any test, analysis, measurement, or monitoring used for the purposes of the Virginia Air Pollution Control Law, the Virginia Waste Management Act or the State Water Control Law (§10.1-1300 et seq., §10.1-1400 et seq., and §62.1-44.2 et seq., respectively, of the Code of Virginia). For the purposes of these regulations, any test, analysis, measurement, or monitoring required [ by pursuant to ] the regulations promulgated under these three laws, or by any permit or order issued under the authority of any of these laws or regulations is "used for the purposes" of these laws. The term shall not include the following:

1. Sampling of [ air, ] water [ or waste, solid and chemical materials, biological tissue, or air and emissions ].

2. Field testing and measurement of [ air, ] water [ or waste, solid and chemical materials, biological tissue, or air and emissions ], except when performed in an environmental laboratory rather than at the site where the sample was taken.

[ 3. Taxonomic identification of samples for which there is no national accreditation standard such as algae, benthic macroinvertebrates, macrophytes, vertebrates and zooplankton. ]

"Environmental laboratory" or "laboratory" means a facility or a defined area within a facility where environmental analysis is performed. [ A structure built solely to shelter field personnel and equipment from inclement weather shall not be considered an environmental laboratory. ]

"Establishment date" means the date set for the accreditation program under this chapter and the certification program under 1VAC30-45 to be established.

"Establishment of accreditation program" or "established program" means that DGS-DCLS has completed the initial accreditation of environmental laboratories covered by this chapter and the initial certification of environmental laboratories covered by 1VAC30-45.

"Facility" means something that is built or installed to serve a particular function.

"Field of accreditation" means an approach to accrediting laboratories by matrix, technology/method and analyte/analyte group.

"Field of accreditation matrix" means the following when accrediting a laboratory:

1. Drinking water. Any aqueous sample that has been designated a potable or potential potable water source.

2. Nonpotable water. Any aqueous sample excluded from the definition of drinking water matrix. Includes surface water, groundwater, effluents, water treatment chemicals, and TCLP or other extracts.

3. Solid and chemical materials. Includes soils, sediments, sludges, products and byproducts of an industrial process that results in a matrix not previously defined.

4. Biological tissue. Any sample of a biological origin such as fish tissue, shellfish, or plant material. Such samples shall be grouped according to origin, i.e., by species.

5. Air and emissions. Whole gas or vapor samples including those contained in flexible or rigid wall containers and the extracted concentrated analytes of interest from a gas or vapor that are collected with a sorbent tube, impinger solution, filter or other device.

"Field of proficiency testing" means an approach to offer proficiency testing by matrix, [ technology technology/method ] , and analyte/analyte group.

"Field testing and measurement" means any of the following:

1. Any test for parameters under 40 CFR Part 136 for which the holding time indicated for the sample requires immediate analysis; or

2. Any test defined as a field test in federal regulation.

The following is a limited list of currently recognized field tests or measures that is not intended to be inclusive: continuous emissions monitoring; on-line monitoring; flow monitoring; tests for pH, residual chlorine, temperature and dissolved oxygen; and field analysis for soil gas.

"Finding" means a conclusion reached during an on-site assessment that identifies a condition having a significant effect on an item or activity. An assessment finding is normally a deficiency and is normally accompanied by specific examples of the observed condition.

"Governmental body" means any department, agency, bureau, authority, or district of the United States government, of the government of the Commonwealth of Virginia, or of any local government within the Commonwealth of Virginia.

"Holding time (or maximum allowable holding time)" means the maximum time that a sample may be held prior to analysis and still be considered valid or not compromised.

"Initial accreditation period" means the period during which DGS-DCLS is accepting and processing applications for the first time under this chapter as specified in 1VAC30-46-70.

"Legal entity" means an entity, other than a natural person, who has sufficient existence in legal contemplation that it can function legally, be sued or sue and make decisions through agents as in the case of corporations.

"Local government" means a municipality (city or town), county, sanitation district, or authority.

[ "Macrophytes" means any aquatic or terrestrial plant species that can be identified and observed with the eye, unaided by magnification. ]

"Matrix" means the component or substrate that contains the analyte of interest.

"National accreditation database" means the publicly accessible database listing the accreditation status of all laboratories participating in NELAP.

"National Environmental Laboratory Accreditation Conference (NELAC)" means a voluntary organization of state and federal environmental officials and interest groups with the primary purpose to establish mutually acceptable standards for accrediting environmental laboratories. A subset of NELAP.

"National Environmental Laboratory Accreditation Program (NELAP)" means the overall National Environmental Laboratory Accreditation Program of which NELAC is a part.

"Noncommercial environmental laboratory" means either of the following:

1. An environmental laboratory where environmental analysis is performed solely for the owner of the laboratory.

2. An environmental laboratory where the only performance of environmental analysis for another person is one of the following:

a. Environmental analysis performed by an environmental laboratory owned by a local government for an owner [ or operator ] of a small [ sewage wastewater ] treatment  [ plant system ] treating domestic sewage at a flow rate of less than or equal to 1,000 gallons per day.

b. Environmental analysis performed by an environmental laboratory operated by a corporation as part of a general contract issued by a local government to operate and maintain a [ sewage wastewater ] treatment [ facility system ] or a waterworks.

c. Environmental analysis performed by an environmental laboratory owned by a corporation as part of the prequalification process [ for or to confirm the identity or characteristics of material supplied by ] a potential [ or existing ] customer [ or generator ] as required by a hazardous waste management permit under 9VAC20-60.

d. Environmental analysis performed by an environmental laboratory owned by a Publicly Owned Treatment Works (POTW) for an industrial source of wastewater under a permit issued by the POTW to the industrial source as part of the requirements of a pretreatment program under Part VII (9VAC25-31-730 et seq.) of 9VAC25-31.

e. Environmental analysis performed by an environmental laboratory owned by a county authority for any municipality within the county's geographic jurisdiction when the environmental analysis pertains solely to the purpose for which the authority was created.

f. Environmental analysis performed by an environmental laboratory owned by an authority or a sanitation district for any participating local government of the authority or sanitation district when the environmental analysis pertains solely to the purpose for which the authority or sanitation district was created.

"Owner" [ or "operator" ] means any person who owns [ or, ] operates [ , leases or controls ] an environmental laboratory.

"Person" means an individual, corporation, partnership, association, company, business, trust, joint venture or other legal entity.

"Physical," for the purposes of fee test categories, means the tests to determine the physical properties of a sample. Tests for solids, turbidity and color are examples of physical tests.

"Pretreatment requirements" means any requirements arising under Part VII (9VAC25-31-730 et seq.) of 9VAC25-31 including the duty to allow or carry out inspections, entry or monitoring activities; any rules, regulations, or orders issued by the owner of a POTW; or any reporting requirements imposed by the owner of a POTW or by the regulations of the State Water Control Board. Pretreatment requirements do not include the requirements of a national pretreatment standard.

"Primary accrediting authority" means the agency or department designated at the territory, state or federal level as the recognized authority with the responsibility and accountability for granting NELAC accreditation to a specific laboratory for a specific field of accreditation.

"Proficiency test or testing (PT)" means evaluating a laboratory's performance under controlled conditions relative to a given set of criteria through analysis of unknown samples provided by an external source.

"Proficiency test (PT) sample" means a sample, the composition of which is unknown to [ both ] the analyst [ and the laboratory ] , provided to test whether the analyst or laboratory or both can produce analytical results within specified acceptance criteria.

"Proficiency testing (PT) program" means the aggregate of providing rigorously controlled and standardized environmental samples to a laboratory for analysis, reporting of results, statistical evaluation of the results and the collective demographics and results summary of all participating laboratories.

"Publicly Owned Treatment Works (POTW)" means a treatment works as defined by §212 of the CWA, which is owned by a state or municipality (as defined by §502(4) of the CWA). This definition includes any devices and systems used in the storage, treatment, recycling, and reclamation of municipal sewage or industrial wastes of a liquid nature. It also includes sewers, pipes, and other conveyances only if they convey wastewater to a POTW treatment plant. The term also means the municipality as defined in §502(4) of the CWA, which has jurisdiction over the indirect discharges to and the discharges from such a treatment works.

"Quality assurance" means an integrated system of activities involving planning, quality control, quality assessment, reporting and quality improvement to ensure that a product or service meets defined standards of quality with a stated level of confidence.

"Quality assurance officer" means the person who has responsibility for the quality system and its implementation. Where staffing is limited, the quality assurance officer may also be the technical director.

"Quality control" means the overall system of technical activities whose purpose is to measure and control the quality of a product or service so that it meets the needs of users.

"Quality manual" means a document stating the management policies, objectives, principles, organizational structure and authority, responsibilities, accountability, and implementation of an agency, organization, or laboratory, to ensure the quality of its product and the utility of its product to its users.

"Quality system" means a structured and documented management system describing the policies, objectives, principles, organizational authority, responsibilities, accountability, and implementation plan of an organization for ensuring quality in its work processes, products (items), and services. The quality system provides the framework for planning, implementing, and assessing work performed by the organization and for carrying out required quality assurance and quality control.

"Quality system matrix," for purposes of batch and quality control requirements, means the following:

1. Aqueous. Any aqueous sample excluded from the definition of drinking water matrix or saline/estuarine source. Includes surface water, groundwater, effluents, and TCLP or other extracts.

2. Drinking water. Any aqueous sample that has been designated a potable or potential potable water source.

3. Saline/estuarine. Any aqueous sample from an ocean or estuary, or other salt water source such as the Great Salt Lake.

4. Non-aqueous liquid. Any organic liquid with less than 15% settleable solids.

5. Biological tissue. Any sample of a biological origin such as fish tissue, shellfish, or plant material. Such samples shall be grouped according to origin [ , i.e., by species ].

6. Solids. Includes soils, sediments, sludges and other matrices with [ less more ]  than 15% settleable solids.

7. Chemical waste. A product or byproduct of an industrial process that results in a matrix not previously defined.

8. Air and emissions. Whole gas or vapor samples including those contained in flexible or rigid wall containers and the extracted concentrated analytes of interest from a gas or vapor that are collected with a sorbent tube, impinger solution, filter or other device.

"Recognition" means the mutual agreement of two or more accrediting authorities to accept each other's findings regarding the ability of environmental laboratories to meet NELAC standards.

"Responsible official" means one of the following, as appropriate:

1. If the laboratory is owned or operated by a private corporation, "responsible official" means (i) a president, secretary, treasurer, or a vice-president of the corporation in charge of a principal business function, or any other person who performs similar policy-making or decision-making functions for the corporation or (ii) the manager of one or more manufacturing, production, or operating facilities employing more than 250 persons or having gross annual sales or expenditures exceeding $25 million (in second-quarter 1980 dollars), if authority to sign documents has been assigned or delegated in accordance with corporate procedures.

2. If the laboratory is owned or operated by a partnership, association, or a sole proprietor, "responsible official" means a general partner, officer of the association, or the proprietor, respectively.

3. If the laboratory is owned or operated by a governmental body, "responsible official" means a director or highest official appointed or designated to oversee the operation and performance of the activities of the governmental laboratory.

4. Any person designated as the responsible official by an individual described in subdivision 1, 2 or 3 of this definition provided the designation is in writing, the designation specifies an individual or position with responsibility for the overall operation of the laboratory, and the designation is submitted to DGS-DCLS.

"Sampling" means the act of collection for the purpose of analysis.

"Sanitation district" means a sanitation district created under the provisions of Chapters 3 (§21-141 et seq.) through 5 (§21-291 et seq.) of Title 21 of the Code of Virginia.

"Sewage" means the water-carried human wastes from residences, buildings, industrial establishments or other places together with such industrial wastes and underground, surface, storm, or other water as may be present.

"Standard operating procedure (SOP)" means a written document which details the method of an operation, analysis or action whose techniques and procedures are thoroughly prescribed and which is accepted as the method for performing certain routine or repetitive tasks.

"TCLP" or "toxicity characteristic leachate procedure" means Test Method 1311 in "Test Methods for Evaluating Solid Waste, Physical/Chemical Methods," EPA Publication SW-846, as incorporated by reference in 40 CFR 260.11. This method is used to determine whether a solid waste exhibits the characteristic of toxicity (see 40 CFR 261.24).

"Technical director (however named)" means the person who has overall responsibility for the technical operation of the environmental laboratory and who exercises actual day-to-day supervision of laboratory operation for the appropriate fields of testing and reporting of results. The title of this person may include but is not limited to laboratory director, technical director, laboratory supervisor or laboratory manager.

"Technology" means a specific arrangement of analytical instruments, detection systems, or preparation techniques, or any combination of these elements.

"Test" means a technical operation that consists of the determination of one or more characteristics or performance of a given product, material, equipment, organism, physical phenomenon, process or service according to a specified procedure.

"Test, analysis, measurement or monitoring required [ by pursuant to ] the Virginia Air Pollution Control Law" means any method of analysis required by the Virginia Air Pollution Control Law (§10.1-1300 et seq.); by the regulations promulgated under this law (9VAC5), including any method of analysis listed either in the definition of "reference method" in 9VAC5-10-20, or listed or adopted by reference in 9VAC5 [ -30, 9VAC5-40, 9VAC5-50 or 9VAC5-60 ]; or by any permit or order issued under and in accordance with this law and these regulations.

"Test, analysis, measurement or monitoring required [ by pursuant to ] the Virginia Waste Management Act" means any method of analysis required by the Virginia Waste Management Act (§10.1-1400 et seq.); by the regulations promulgated under this law (9VAC20), including any method of analysis listed or adopted by reference in 9VAC20 [ -60, 9VAC20-80, 9VAC20-101, or 9VAC20-120 ]; or by any permit or order issued under and in accordance with this law and these regulations.

"Test, analysis, measurement or monitoring required [ by pursuant to ] the Virginia Water Control Law" means any method of analysis required by the Virginia Water Control Law (§62.1-44.2 et seq.); by the regulations promulgated under this law (9VAC25), including any method of analysis listed or adopted by reference in 9VAC25 [ -31, 9VAC25-32, 9VAC25-110, 9VAC25-120, 9VAC25-151, 9VAC25-180, 9VAC25-190, 9VAC25-192, or 9VAC25-210 ] ; or by any permit or order issued under and in accordance with this law and these regulations.

"Test method" means an adoption of a scientific technique for [ performing ] a specific measurement [ problem ] , as documented in a laboratory standard operating procedure or [ as ] published by a recognized authority.

"U.S. Environmental Protection Agency (U.S. EPA or EPA)" means the federal government agency with responsibility for protecting, safeguarding and improving the natural environment (i.e., air, water and land) upon which human life depends.

"Virginia Air Pollution Control Law" means [ Chapter 13 ] §10.1-1300 [ et seq. ] of the Code of Virginia which is titled "Air Pollution Control Board."

[ "Wastewater" means liquid and water-carried industrial wastes and domestic sewage from residential dwellings, commercial buildings, industrial and manufacturing facilities and institutions. ]

"Waterworks" means each system of structures and appliances used in connection with the collection, storage, purification, and treatment of water for drinking or domestic use and the distribution thereof to the public, except distribution piping.

[ "Zooplankton" means microscopic animals that float freely with voluntary movement in a body of water. ]

1VAC30-46-50. Scope of accreditation.

A. Commercial environmental laboratories shall be accredited based on the general laboratory standards set out in Part II (9VAC30-46-200 1VAC30-46-200 ] et seq.) of this chapter and on the specific test methods or analysis, monitoring or measurement required by Virginia Air Pollution Control Law, Virginia Waste Management Act or Virginia Water Control Law, the regulations promulgated under these laws, and by permits and orders issued under and in accordance with these laws and regulations.

B. DGS-DCLS shall review alternative test methods and procedures for accreditation when these are proposed by the applicant laboratory. The provisions of 1VAC30-46-70 E and 1VAC30-46-90 B govern alternative test methods and procedures.

C. Accreditation shall be granted for [ a specific field one ] or [ more ] fields of accreditation, including the [ matrix, the ] technology and methods used by the commercial environmental laboratory, and the individual analytes or analyte groups determined by the particular method.

1VAC30-46-60. General: accreditation requirements.

A. Components of accreditation.  The components of accreditation include review of personnel qualifications, on-site assessment, proficiency testing and quality assurance and quality control standards.  The criteria for these components, specified in Part II ( [ 9VAC30-46-200 1VAC30-46-200 ] et seq.) of this chapter, shall be fulfilled for accreditation.

B. Individual laboratory sites and mobile laboratories.

1. Individual laboratory sites are subject to the same application process, assessments, and other requirements as environmental laboratories.  Any remote laboratory sites are considered separate sites and subject to separate on-site assessments.

2. Laboratories located at the same physical location shall be considered an individual laboratory site if these laboratories are owned [ or operated ] by the same person, and have the same technical director and quality system.

3. [ Laboratories located at separate, noncontiguous physical locations may request to be considered as an individual laboratory site if these laboratories are owned by the same person and have the same laboratory manager and quality system.

4. ] A mobile laboratory, which is configured with equipment to perform environmental analyses, whether associated with a fixed-based laboratory or not, is considered an environmental laboratory and shall require separate accreditation.  This accreditation shall remain with the mobile laboratory and be site independent.  Moving the configured mobile laboratory to a different site shall not require a new or separate accreditation.  Before performing analyses at each new site, the laboratory shall ensure that instruments and equipment have been checked for performance and have been calibrated.

1VAC30-46-70. Process to apply and obtain accreditation.

A. Duty to apply. All owners [ or operators ] of (i) commercial environmental laboratories and (ii) [ NELAC NELAP ] -accredited environmental laboratories [ located outside Virginia ] applying for reciprocal accreditation shall apply for accreditation as specified by the provisions of this section.

B. Timely initial applications.

1. Owners [ or operators ] of commercial environmental laboratories applying for accreditation under this chapter for the first time shall submit an application  to DGS-DCLS no later than [ 180 calendar days after the effective date of this chapter March 30, 2009 ].

2. Owners [ or operators ] of commercial environmental laboratories that come into existence after [ this chapter becomes effective October 1, 2008, ] shall submit an initial application to DGS-DCLS no later than 180 calendar days prior to initiating the provision of environmental laboratory services.

3. Owners [ or operators ] of [ NELAC NELAP ] -accredited environmental laboratories [ located outside Virginia ].

a. During the initial accreditation period, [ NELAC NELAP ] -accredited environmental laboratories [ located outside Virginia ] shall submit an application to DGS-DCLS no later than [ 180 calendar days after the effective date of this chapter March 30, 2009 ].

b. After the program is established, [ NELAC NELAP ] -accredited environmental laboratories [ located outside Virginia ] shall submit an application to DGS-DCLS no later than 180 calendar days prior to initiating the provision of environmental laboratory services [ in Virginia ].

C. Timely renewal applications. [ The owner or operator of either an (i) accredited commercial environmental laboratory or (ii) environmental laboratory holding reciprocal accreditation shall submit an application for renewal of accreditation at least 90 calendar days prior to expiration of accreditation.

1. Every two years from the date of initial accreditation, laboratories accredited under this chapter shall submit an application for renewal of accreditation as required by subsection F of this section, including the fees required by 1VAC30-46-150. During this biannual renewal DGS-DCLS shall perform an on-site assessment in addition to a review of the laboratory’s application package.

2. Every other year, DGS-DCLS shall renew accreditation for the accredited laboratory provided the laboratory does all of the following:

a. Maintains compliance with this chapter.

b. Attests to this compliance by signing the Certificate of Compliance provided under subdivision F 3 of this section.

c. Reports acceptable proficiency test values for the Fields of Accreditation for which the laboratory held accreditation during the previous year.

The laboratory shall submit the application information required by subdivisions F 1 (except for the quality manual) and F 3 of this section.

3. Renewal application due dates.

a. The owner of either an (i) accredited commercial environmental laboratory or (ii) environmental laboratory holding reciprocal accreditation shall submit an application for renewal of accreditation under subdivision C 1 of this section at least 90 calendar days prior to expiration of accreditation.

b. The owner of either an (i) accredited commercial environmental laboratory or (ii) environmental laboratory holding reciprocal accreditation shall submit an application for renewal of accreditation under subdivision C 2 of this section at least 30 calendar days prior to expiration of accreditation. ]

D. Responsibilities of the owner [ or and ] operator [ when the laboratory is owned by one person and operated by another person ].

1. When an environmental laboratory is owned by one person but is operated by another person, the operator may submit the application for the owner.

2. If the operator fails to submit the application, the owner is not relieved of his responsibility to apply for accreditation.

3. While DGS-DCLS may notify environmental laboratories of the date their applications are due, failure of DGS-DCLS to notify does not relieve the owner [ or operator ] of his obligation to apply under this chapter.

E. Submission of applications for modifications to accreditation. An owner [ or operator ] of an accredited environmental laboratory shall follow the process set out in 1VAC30-46-90 B to add a new [ technology matrix technology/method ] , an analyte or [ a test method, analyte group ], modify a test method or institute use of a method [ or technology ] not in the laboratory’s standard operating procedures, including alternative test methods or procedures.

F. Contents of application.

1. Applications shall include the following information and documents:

a. Legal name of laboratory;

b. Name of owner of laboratory;

c. Name of operator of laboratory, if different than owner;

d. Street address and description of location of laboratory;

e. Mailing address of laboratory, if different from street address;

f. Address of owner, if different from laboratory address;

g. Name, address, telephone number, facsimile number and e-mail, as applicable, of responsible official;

h. Name, address, telephone number, facsimile number and e-mail, as applicable, of technical director;

i. Name, address, telephone number, facsimile number and e-mail, as applicable, of designated quality assurance officer;

j. Name [ , title ] and telephone number of laboratory contact person;

k. Laboratory type (e.g., commercial, public wastewater system, mobile);

l. Laboratory hours of operation;

m. Fields of accreditation for which the laboratory is seeking accreditation;

n. Methods employed, including analytes;

o. The results of the three most recent proficiency test studies;

p. Quality assurance manual;

q. Lab identification number (for renewal only); and

r. For mobile laboratories, a unique vehicle identification number, such as a manufacturer’s vehicle identification number (VIN #), serial number, or license number.

2. Fee. The application shall include payment of the fee as specified in 1VAC30-46-150.

3. Certification of compliance.

a. The application shall include a "Certification of Compliance" statement signed and dated by the quality assurance officer, and a responsible official or the technical director or both.

b. The certification of compliance shall state: "The applicant understands and acknowledges that the laboratory is required to be continually in compliance with the Virginia environmental laboratory accreditation program regulation (1VAC30 Chapter 46) and is subject to the provisions of 1VAC30-46-100 in the event of noncompliance. I certify under penalty of law that this document and all attachments were prepared under my direction or supervision in accordance with a system designed to assure that qualified personnel properly gather and evaluate the information submitted. Based on my inquiry of the person or persons who manage the laboratory or those persons directly responsible for gathering and evaluating the information, the information submitted is, to the best of my knowledge and belief, true, accurate and complete. Submitting false information or data shall result in denial or withdrawal of accreditation. I further certify that I am authorized to sign this application."

G. Completeness determination.

1. DGS-DCLS shall determine whether an application is complete and notify the laboratory of the result of such determination. [ Except during During ] the initial accreditation period, DGS-DCLS shall provide this notice within [ 60 90 ] calendar days of [ DGS-DCLS's its ] receipt of the [ initial ] application. [ Following the initial accreditation period, DGS-DCLS shall provide this notice as follows:

a. Within 60 calendar days of DGS-DCLS' receipt of a laboratory's initial application.

b. Within 30 calendar days of DGS-DCLS’ receipt of a laboratory’s renewal application under subdivision C 1 of this section.

c. Within 15 calendar days of DGS-DCLS’ receipt of a laboratory’s renewal application under subdivision C 2 of this section. ]

2. An [ initial ] application [ or an application for renewal under subdivision C 1 of this section ] shall be determined complete if it contains all the information required pursuant to subsection F of this section and is sufficient to evaluate the laboratory prior to the on-site assessment. [ DGS-DCLS shall consider an application for renewal under subdivision C 2 of this section to be complete if it contains the information required under subdivision C 2 of this section ] Designating an application complete does not preclude DGS-DCLS from requesting or accepting additional information.

3. If DGS-DCLS determines that an application is incomplete, DGS-DCLS's notification of such determination shall explain why the application is incomplete and specify the additional information needed to make the application complete.

4. Except during the initial accreditation period, if no determination is made within 60 calendar days of DGS-DCLS's receipt of either (i) the application or (ii) additional information, in the case of an application determined to be incomplete, the application shall be determined to be complete. [ During the initial accreditation period, the time period shall be 90 calendar days. ]

5.  [ If the laboratory has not submitted the required additional information within 90 days of receiving a notice from ] DGS-DCLS [ may deny any application from a laboratory and require the laboratory to submit a new application if the laboratory does not submit additional information required by DGS-DCLS within 90 days of the mailing date of the notice that requires additional information requesting additional information, DGS-DCLS may return the incomplete application and inform the laboratory that the application cannot be processed. The laboratory may then submit a new application ].

H. Grant of interim accreditation pending final determination on application.

1. DGS-DCLS shall grant [ a laboratory ] interim accreditation status [ to laboratories applying initially or for renewal under subdivision C 1 of this section ] under the following conditions:

a. The laboratory’s application is determined to be complete;

b. The laboratory has satisfied all the requirements for accreditation, including all requests for additional information, with the exception of on-site assessment; and

c. DGS-DCLS is unable to schedule the on-site assessment within 90 days of its determination that the application is complete [ and that the laboratory has satisfied all other requirements for accreditation (for initial applications) or before the laboratory's accreditation expires (for renewal applications under subdivision C 1 of this section) ].

[ 2. DGS-DCLS shall grant interim accreditation status to a laboratory renewing its accreditation under subdivision C 2 of this section during its review of the renewal application if the owner has submitted a complete application as required under subdivision C 2 of this section. ]

[ 2. 3. ]  A laboratory with interim accreditation [ status ] shall have the same rights and status as a laboratory that has been granted accreditation by DGS-DCLS.

[ 3. 4. ] Interim accreditation [ expires when DGS-DCLS issues a final determination on accreditation status shall not exceed 12 months ].

I. On-site assessment.

[ 1. ] An on-site assessment shall be performed and the follow-up and reporting procedures for such assessments shall be completed in accordance with Part II [ (9VAC30-46-200 (1VAC30-46-200 ] et seq.) of this chapter prior to issuance of a final determination on accreditation.

[ 2. Alternative on-site assessment option. If DGS-DCLS is unable to schedule an on-site assessment under the conditions of subdivision H 1 c of this section, the owner of the applicant laboratory may use third-party on-site assessors instead of DGS-DCLS on-site assessors under the following conditions:

a. The third-party on-site assessors are on a DGS-DCLS-approved list of NELAC-trained on-site assessors, and

b. The owner of the applicant laboratory agrees to pay the third-party on-site assessors. ]

J. Final determination on accreditation.

1. Upon completion of the accreditation review process and corrective action, if any, DGS-DCLS shall grant accreditation in accordance with subsection K of this section or deny accreditation in accordance with subsection L of this section.

2. Except during the initial accreditation period, DGS-DCLS shall complete action on a laboratory’s application within nine months from the time [ an application is determined to be complete a completed application is received from the laboratory ].

3. During the initial accreditation period, DGS-DCLS shall notify applicants of their interim accreditation status under subsection H of this section only after all applications have been reviewed and are determined to be complete.

4. During the final approval process of the initial accreditation period, DGS-DCLS shall notify applicants of their final accreditation status only after all timely and complete applications have been reviewed, all on-site assessments have been completed, and accreditation status has been determined for all applicant laboratories.

5. During the final approval process, DGS-DCLS shall release on-site assessment reports to applicants at the time that applicants are notified of their final accreditation status. [ If a laboratory is found to have deficiencies during the on-site assessment, DGS-DCLS may provide comments and recommendations aimed at helping the laboratory improve. ]

K. Grant of accreditation.

1. When a laboratory meets the requirements specified for receiving accreditation, DGS-DCLS shall issue a certificate to the laboratory.  The certificate shall be sent to the technical director, and the responsible official shall be notified.

2. The [ certificate shall be signed by the ] director of DGS-DCLS [ shall sign the certificate ]. The certificate shall be transmitted as a sealed and dated document.

3. The certificate shall include the following information:

a. Name of owner [ or operator ] of laboratory;

[ b. Name of operator of laboratory, if different from owner; ]

[ b. c. ] Name of responsible official;

[ c. d. ] Address and location of laboratory;

[ d. e. ] Laboratory identification number;

[ e. f. ] Fields of accreditation (matrix, technology/method and analyte/analyte group) for which accreditation is granted;

[ f. g. ]  Any addenda or attachments; and

[ g. h. ] Issuance date and expiration date.

4. National Environmental Laboratory Accreditation Program (NELAP) status.

a. Laboratories accredited under this chapter are accredited under the standards of the National Environmental Laboratory Accreditation Conference.

b. The certificate of accreditation shall contain the NELAP insignia.

c. Accredited laboratories shall comply with the provisions of 1VAC30-46-130 with regard to the use of these certificates and their status as NELAP-accredited laboratories.

5. The laboratory shall post the most recent certificate of accreditation and any addenda to the certificate issued by DGS-DCLS in a prominent place in the laboratory facility.

6. Accreditation shall expire [ two years one year ] after the date on which accreditation is granted.

L. Denial of accreditation.

1. DGS-DCLS shall deny accreditation to an environmental laboratory in total if the laboratory [ owner or an employee falsifies is found to be falsifying ] any data or [ provides providing ] false information to support accreditation.

2. Denial of accreditation in total or in part.

a. DGS-DCLS may deny accreditation to an environmental laboratory in total or in part if the laboratory [ owner or an employee ] fails to do any of the following:

(1) Pay the required fees;

(2) Employ laboratory staff to meet the personnel qualifications as required by Part II (1VAC30-46-200 et seq.) of this chapter;

(3) Successfully analyze and report proficiency testing samples as required by Part II of this chapter;

(4) Submit a corrective action report in accordance with Part II of this chapter in response to a deficiency report from the on-site assessment team within the required 30 calendar days;

(5) Implement the corrective actions detailed in the corrective action report within the time frame specified by DGS-DCLS;

(6) Pass required on-site assessment as specified in Part II of this chapter;

(7) Implement a quality system as defined in Part II of this chapter.

b. DGS-DCLS may deny accreditation to an environmental laboratory in total or in part if the laboratory’s application is not determined to be complete within 90 days following notification of incompleteness because the laboratory is delinquent in submitting information required by DGS-DCLS in accordance with this chapter.

c. DGS-DCLS may deny accreditation to an environmental laboratory in total or in part if the DGS-DCLS on-site assessment team is unable to carry out the on-site assessment pursuant to 1VAC30-46-210 B because [ an employee, owner,  or other a ] representative of the environmental laboratory denied the team entry during [ the laboratory's ] normal business hours [ that it specified in the laboratory application ].

3. [ To deny accreditation, DGS-DCLS shall provide by certified mail written notification of denial to the responsible officer and the technical director of the laboratory, including a detailed explanation of the reason for denial and notice of the right to appeal such denial. DGS-DCLS shall follow the process specified in 1VAC30-46-110 when denying accreditation to an environmental laboratory. ]

M. Reapplication following denial of accreditation.

1. Upon denial of accreditation, the laboratory shall wait six months before reapplying for accreditation.

2. DGS-DCLS shall not waive application fees for a laboratory reapplying for accreditation.

1VAC30-46-80. Maintaining accreditation.

A. Accreditation remains in effect until withdrawn by DGS-DCLS, withdrawn voluntarily at the written request of the accredited laboratory, or expiration of the accreditation period.  To maintain accreditation, the accredited laboratory shall comply with the elements listed in this section and in 1VAC30-46-90.

B. Quality systems.  A laboratory seeking to maintain accreditation under this regulation shall assure consistency and promote the use of quality assurance and quality control procedures.  Part II ( [ 9VAC30-46-200 1VAC30-46-200 ] et seq.) of this chapter specifies the quality assurance and quality control requirements that shall be met to maintain accreditation.  The laboratory shall establish and maintain a quality system based on the required elements contained in Part II [ and appropriate to the type, range and volume of environmental testing activities it undertakes ].

C. Proficiency tests.  Laboratories seeking to maintain accreditation under this regulation shall perform proficiency tests as required under Part II ( [ 9VAC30-46-200 1VAC30-46-200 ] et seq.) of this chapter.

D. Recordkeeping and retention.  All laboratory records associated with accreditation parameters shall be kept as provided by the requirements for records under Part II [ 9VAC30-46-200 1VAC30-46-200 ] et seq.) of this chapter.  These records shall be maintained for a minimum of five years unless designated for a longer period by another regulation or authority.  All such records shall be available to DGS-DCLS upon request.

1VAC30-46-90. [ Changing Notifications and changes to ] accreditation [ elements and ] status.

A. Changes to key accreditation criteria. [ 1. ] The accredited laboratory shall notify DGS-DCLS [ as set out in subdivision 2 of this subsection in writing ] of any changes in key accreditation criteria within 30 calendar days of the change. Key accreditation criteria are laboratory ownership, location, key personnel, [ test methods, analytes, ] and major instrumentation.

[ 2. The laboratory may initially notify DGS-DCLS of any change to key accreditation criteria by e-mail, facsimile or telephone.  The notification by e-mail, facsimile or telephone subsequently shall be submitted in writing.

3. As specified in subsection B of this section, changes to key accreditation criteria that affect the laboratory’s scope of accreditation require review and approval by DGS-DCLS in advance of the laboratory’s making the change. ]

B. Changes to scope of accreditation.

1. DGS-DCLS [ shall review and may ] approve [ the addition of a laboratory's application to add ] a new [ matrix, ] technology, [ an ] analyte, or [ a ] test method to a laboratory’s scope of accreditation [ or otherwise modify the laboratory's scope of accreditation by performing a data review ].

2. To [ begin the process of review apply ] , the owner [ or operator ] of the accredited laboratory [ that wants to add to the laboratory’s scope of accreditation ] shall submit the following [ application materials ] to DGS-DCLS:

a. A letter signed by the owner [ or operator ] that briefly summarizes the addition to be made to the laboratory’s scope of accreditation.

b. Pertinent information demonstrating [ that the laboratory is capable of performing the test method or using the technology to be added the laboratory's capability to perform the additional matrix, technology/method, or analyte/analyte group, ] such as proficiency testing performance and quality control performance.

c. A written standard operating procedure covering the new method, analyte, or technology [ matrix, technology/method or analyte/analyte group ].

[ DGS-DCLS may request additional material to complete its review. ]

3. DGS-DCLS may approve a laboratory’s application for modification to its scope of accreditation by performing a review of the application materials submitted, without an on-site assessment. [ An The ] addition of a new technology or test method requiring [ the use of ] specific equipment may require an on-site assessment.  Other reviews of performance and documentation may be carried out by DGS-DCLS, depending on the modification for which the laboratory applies.

4. [ Within 90 calendar days of the receipt of the application from the accredited environmental laboratory, DGS-DCLS shall review and determine whether the proposed modification may be approved.

5. ] If the proposed modification to the laboratory’s scope of accreditation is approved, DGS-DCLS shall amend the laboratory’s certificate of accreditation.

C. Change of ownership or location of laboratory.

1. The accredited laboratory shall submit a written notification to DGS-DCLS of the change of ownership or location of the laboratory within 30 calendar days of the change. [ This requirement applies only to fixed-based and not mobile laboratories. ]

2. Accreditation may be transferred when the legal status or ownership of a accredited laboratory changes [ without affecting its as long as the transfer does not affect the laboratory's ] personnel, equipment, [ and facilities or organization ].

3. [ DGS-DCLS may charge a transfer fee and may conduct an on-site assessment to verify the effects of such changes on laboratory performance. If the laboratory’s personnel, equipment, or organization are affected by the change of legal status or ownership, DGS-DCLS may require reaccreditation or reapplication in any or all of the categories for which the laboratory is accredited.

4. DGS-DCLS may require an on-site assessment depending on the nature of the change of legal status or ownership. DGS-DCLS shall determine the elements of any on-site assessment required. ]

[ 4. 5. ] When [ a laboratory changes ] there is a change in ] ownership, the new [ laboratory ] owner [ of the accredited laboratory ] shall assure [ that the history of the laboratory's ownership can be traced through laboratory identification numbers historical traceability of the laboratory accreditation numbers. ]

[ 5. 6. ] When there is a change in ownership, [ the new owner of the accredited laboratory shall keep ] all records and analyses performed by the previous owner under his scope of accreditation [ shall be kept ] for a period of five years. [ As required under 1VAC30-46-80 D, all such records shall be made available to DGS-DCLS upon request. These records and analyses are subject to inspection by DGS-DCLS during this five-year period. This provision applies regardless of change of ownership, accountability or liability. ]

D. Voluntary withdrawal.  Any environmental laboratory owner [ or operator ] who wishes to withdraw the laboratory from its accreditation status or from being accredited, in total or in part, shall submit written notification to DGS-DCLS no later than 30 calendar days before the end of the laboratory’s accreditation term.  Within 30 calendar days, DGS-DCLS shall provide the laboratory with a written notice of withdrawal.

1VAC30-46-100. Withdrawal of accreditation.

A. DGS-DCLS shall withdraw accreditation from an environmental laboratory in total [ for the following reasons if the laboratory is found to be falsifying any data or providing false information to support certification ]:

[ 1. Submittal by the laboratory owner or employee of proficiency test sample results generated by another laboratory as its own.

2. Falsification by a laboratory owner or employee of any data or the provision of false information by any laboratory owner or employee to support accreditation.

3. Conviction of the laboratory owner or employee of charges relating to the falsification of any report concerning a laboratory analysis. ]

B. DGS-DCLS may withdraw accreditation from an environmental laboratory in part or in total when the laboratory [ owner or an employee ] has failed to do any of the following:

1. Participate in the proficiency testing program as required by 1VAC30-46-210 C.

2. Complete proficiency testing studies and maintain a history of at least two successful proficiency testing studies for each affected accredited field of testing out of the three most recent proficiency testing studies as defined in 1VAC30-46-210 C.

3. Maintain a quality system as defined in 1VAC30-46-210 D.

4. Employ staff that meet the personnel qualifications of 1VAC30-46-210 A.

5. Submit an acceptable corrective action report after two opportunities as specified in 1VAC30-46-210 B.

6. Implement corrective action specified in the laboratory’s corrective action report as set out under 1VAC30-46-210 B.

7. Notify DGS-DCLS of any changes in key accreditation criteria as set forth in 1VAC30-46-90.

8. Use correct and authorized references to the laboratory’s accreditation status or that of DGS-DCLS in the laboratory’s documentation and advertising as set forth in 1VAC30-46-130.

C. [ DGS-DCLS shall follow the process specified in 1VAC30-46-110 when withdrawing accreditation from an environmental laboratory.

D. ] Responsibilities of the environmental laboratory and DGS-DCLS when accreditation has been withdrawn.

1. Laboratories that lose their accreditation in full shall return their certificate to DGS-DCLS.

2. If a laboratory loses accreditation in part, an addendum to the certificate shall be issued by DGS-DCLS to the laboratory.

3. The laboratory shall discontinue the use of all materials that contain either a reference to the environmental laboratory’s past accreditation status or that display the NELAC/NELAP logo. These materials may include catalogs, advertising, business solicitations, proposals, quotations, laboratory analytical results or other materials.

[ D. E. ] After correcting the reason or cause for the withdrawal of accreditation under 1VAC30-46-100 A or B, the laboratory owner [ or operator ] may reapply for accreditation.

1VAC30-46-110. [ Appeal  Procedures to deny accreditation, to withdraw accreditation, and appeal ] procedures.

A. [ If DGS-DCLS believes it has grounds to deny accreditation to or withdraw accreditation from an environmental laboratory, ] DGS-DCLS shall notify [ an the ] environmental laboratory in writing of its [ decision to deny accreditation to or to withdraw accreditation from an environmental laboratory intent to hold an informal fact finding under §2.2-4019 of the Code of Virginia in order to make a decision on the denial of accreditation or withdrawal of accreditation. DGS-DCLS shall send this notification by certified mail to the responsible official and provide a copy to the technical director of the environmental laboratory. The notice of informal fact finding shall provide a detailed explanation of the basis for the notice ].

B. [ All appeals taken from actions of the DGS-DCLS director relative to the provisions of this chapter shall be governed by the Virginia Administrative Process Act (§2.2-4000 et seq. of the Code of Virginia). Following the informal fact finding held pursuant to §2.2-4019 of the Code of Virginia, the director shall render a decision regarding accreditation, and shall send this notification by certified mail to the responsible official and provide a copy to the technical director of the environmental laboratory. If the director’s decision is adverse to the environmental laboratory, the responsible official may appeal this decision in accordance with §2.2-4026 of the Code of Virginia and Part 2A of the Rules of the Supreme Court of Virginia.

C. The provisions of this section do not preclude informal discussions between DGS-DCLS and any environmental laboratory that has been notified of a possible denial or withdrawal of accreditation. These informal discussions to resolve the concerns that prompted the notice shall be held prior to the informal fact finding proceeding.

D. The accreditation status of an environmental laboratory appealing withdrawal of accreditation shall not change pending the final decision of the appeals filed under the Virginia Administrative Process Act (§2.2-4000 et seq. of the Code of Virginia) and Part 2A of the Rules of the Supreme Court of Virginia. ]

1VAC30-46-120. National accreditation database.

DGS-DCLS shall provide to NELAP the following information about environmental laboratories accredited under this chapter: (i) technical director’s name; (ii) ownership and location of laboratory and any changes; (iii)  key accreditation criteria and any changes; (iv) interim, as well as final, accreditation status; and (v) on-site assessment reports.

1VAC30-46-130. Use of accreditation status by environmental laboratories accredited under this chapter.

A. The owner [ or operator ] of an environmental laboratory accredited under this chapter shall not misrepresent the laboratory’s fields of accreditation or its accreditation status on any document. This includes laboratory reports, catalogs, advertising, business solicitations, proposals, quotations or other materials.

B. Environmental laboratories accredited under this chapter shall comply with all of the following:

1. Post or display their most recent accreditation certificate or their fields of accreditation in a prominent place in the laboratory facility.

2. Make accurate statements concerning their fields of accreditation and accreditation status.

3. Accompany DGS-DCLS’s name or the NELAC/NELAP logo or both with at least the phrase “NELAP-accredited” and the laboratory’s identification number or other identifier when DGS-DCLS’s name is used on general literature such as catalogs, advertising, business solicitations, proposals, quotations, laboratory analytical reports or other materials.

4. Not use their accreditation certificate, their accreditation status or the NELAC/NELAP logo to imply endorsement by DGS-DCLS.

C. The owners [ or operators ] of laboratories accredited under this chapter who choose to (i) use DGS-DCLS’s name; (ii) make reference to its NELAP accreditation status; or (iii) use the NELAC/NELAP logo in any catalogs, advertising, business solicitations, proposals, quotations, laboratory analytical reports or other materials, shall comply with both of the following:

1. Distinguish between proposed testing for which the laboratory is accredited and the proposed testing for which the laboratory is not accredited.

2. Include the laboratory’s identification number or other identifier.

1VAC30-46-140. Reciprocal accreditation.

A. DGS-DCLS, when recognized by NELAP as a primary accrediting authority, may grant reciprocal accreditation to an environmental laboratory [ located outside Virginia ] that holds a current accreditation from another NELAP-recognized primary accrediting authority.

B. The owner [ or operator ] of a NELAP-accredited environmental laboratory that seeks accreditation under this chapter shall apply as specified in 1VAC30-46-70.

C. The owner [ or operator ] of the applicant laboratory shall pay the fee required by 1VAC30-46-150.

D. DGS-DCLS shall not require a NELAP-accredited environmental laboratory that seeks accreditation under this section to meet any additional proficiency testing, quality assurance, or on-site assessment requirements for the fields of accreditation for which the laboratory holds primary NELAP accreditation.

E. DGS-DCLS shall consider only the current certificate of accreditation issued by the NELAP-recognized primary accrediting authority.

F. DGS-DCLS shall do the following:

1. Grant reciprocal accreditation for only the fields of accreditation for which the laboratory holds current primary NELAP accreditation.

2. Except during the initial accreditation period, grant reciprocal accreditation and issue certificates to an applicant laboratory within 30 calendar days of receipt of the laboratory's application.

G. Potential nonconformance issues.

1. If DGS-DCLS notes any potential nonconformance with the NELAC standards by a laboratory during the initial application process for reciprocal accreditation or for a laboratory that already has been granted NELAP accreditation through reciprocal accreditation, DGS-DCLS shall immediately notify, in writing, the applicable NELAP-recognized primary accrediting authority and the laboratory. The notification shall cite the applicable sections within the NELAC standards for which nonconformance by the laboratory has been noted.

2. If the alleged nonconformance is noted during the initial application process for reciprocal accreditation, final action on the application for reciprocal accreditation shall not be taken until the alleged nonconformance issue has been resolved.

3. If the alleged nonconformance is noted after reciprocal accreditation has been granted, the laboratory shall maintain its current accreditation status until the alleged nonconformance issue has been resolved.

4. If DGS-DCLS does not believe the primary accrediting authority has taken timely and appropriate action on the potential nonconformance, DGS-DCLS shall notify the NELAP director of its concerns.

1VAC30-46-150. Fees.

A. General.

1. Fees shall be submitted with all applications [ , including reapplications, ] for accreditation [ and all renewal applications for accreditation under 1VAC30-46-70 C 1 ]. Applications shall not be designated as complete until the fee is received by DGS-DCLS.

2. Fees shall be nonrefundable.

3. An environmental laboratory applying for reciprocal accreditation under this chapter shall pay the same fee as other laboratories subject to this chapter.

B. Fee computation.

1. The fee shall be the total of the base fee and the test category fees.

2. The test category fees cover categories for the test methods to be accredited as specified in the laboratory’s application.

3. If the total of the base fee and the test category fees is more than the maximum fee, the laboratory shall pay the maximum fee.

[ C. Maximum fee. The maximum fee shall be $4,200.

D. C. ]  Base fee. The base fee shall be [ $2,100 $1,700 ].

[ D. Maximum fee. The maximum fee shall be $5,200. ]

E. Test category fees.

1. Fees shall be charged for each category of tests to be accredited.

2. The fee for each category includes one or more analytical methods unless otherwise specified. With the exception of the test categories labeled oxygen demand and physical, test categories related to test methods for water are defined by 40 CFR 136.3.

3. Fees.

|  | TEST CATEGORY | FEE |
| --- | --- | --- |
|  | Oxygen demand (BOD or COD) | [ $300 $375 ] |
|  | Bacteriology | [ $300 $375 ] |
|  | Inorganic chemistry, fewer than four methods | [ $300 $375 ] |
|  | Inorganic chemistry, four or more methods | [ $600 $750 ] |
|  | Chemistry metals, [ fewer than four one - two ] methods | [ $300 $450 ] |
|  | Chemistry metals, [ four or ] more [ than two ] methods | [ $600 $1,000 ] |
|  | Organic chemistry, fewer than four methods | [ $350 $600 ] |
|  | Organic chemistry, four or more [ methods ] | [ $700 $1,200 ] |
|  | [ Whole effluent Aquatic ] toxicity, acute methods only | [ $300 $400 ] |
|  | [ Whole effluent Aquatic ] toxicity, acute and chronic methods | [ $600 $700 ] |
|  | Radiochemical | [ $900 $1,000] |
|  | Physical | [ $300 $375 ] |

F. Additional fees. [ Additional fees shall be charged to laboratories applying for the following: (i) modification to scope of accreditation under 1VAC30-46-90 B, (ii) transfer of ownership under 1VAC30-46-90 C, (iii) request that multiple, noncontiguous laboratory sites be considered as one site under 1VAC30-46-60 B 3, or (iv) petition for a variance under 1VAC30-46-160. ]

1. For any accredited environmental laboratory that applies to modify its scope of accreditation as specified under 1VAC30-46-90 B, DGS-DCLS shall assess a fee determined by the method in [ subdivision 3 subsection G ] of this [ subsection section ] .

2. Under 1VAC30-46-90 C, DGS-DCLS may charge a transfer fee to a certified laboratory that transfers ownership. [ If DGS-DCLS determines that an on-site assessment is necessary to evaluate the effect of the transfer of ownership, DGS-DCLS shall assess a fee determined by the method in subdivision 3 of this subsection. A fee shall be charged if DGS-DCLS (i) needs to review documentation sent by the laboratory about the transfer of ownership or (ii) determines that an on-site assessment is necessary to evaluate the effect of the transfer of ownership. DGS-DCLS shall assess a fee determined by the method in subsection G of this section. ] If DGS-DCLS determines that a fee should be charged, the fee shall be a minimum of $100 and a maximum of $1,000. [ If, under 1VAC30-46-90 C, DGS-DCLS determines that the change of ownership or location of laboratory requires reaccreditation of or reapplication by the laboratory, the laboratory shall pay the application fee required under this section.

3. Under 1VAC30-46-60 B 3, the owner of multiple non-contiguous laboratories may request that DGS-DCLS consider these laboratories to be one site. If, as a result of the request being granted, DGS-DCLS needs to perform multiple on-site assessments, DGS-DCLS shall charge a fee for the additional on-site assessments. The fee shall be the sum of reasonable travel costs and labor charges for the additional on-site assessments. The labor charges will be determined following the method in subsection G of this section.

4. Under 1VAC30-46-160, any person regulated by this chapter may petition the director to grant a variance from any requirement of this chapter. DGS-DCLS shall charge a fee for the time needed to review the petition, including any on-site assessment required. The fee shall be determined by the method specified in subsection G of this section. ]

[ 3. G. ] Fee determination.

[ a. 1. ] The fee shall be the sum of the total hourly charges for all reviewers plus any on-site review costs incurred.

[ b. 2. ] An hourly charge per reviewer shall be determined by (i) obtaining a yearly cost by multiplying the reviewer’s annual salary by 1.35 (accounts for overhead such as taxes and insurance) and then (ii) dividing the yearly cost by 1,642 (number of annual hours established by [ DGS fiscal services Fiscal Services, DGS, ] for billing purposes).

[ c. 3. ]  The charge per reviewer shall be determined by multiplying the number of hours expended in the review by the reviewer’s hourly charge.

[ d. 4. ] If an on-site review is required, travel time and on-site review time shall be charged at the same hourly charge per reviewer, and any travel expenses shall be added.

[ G. On-site assessment fees. When, with the concurrence of the applicant laboratory, DGS-DCLS uses approved, third-party on-site assessors, the cost of the on-site assessment shall be paid by the applicant. H. Out-of-state laboratories – travel costs. The owner of an environmental laboratory located in another state who applies for accreditation under this chapter shall also pay a fee equal to the reasonable travel costs associated with conducting an on-site assessment at the laboratory. Reasonable travel costs include transportation, lodging, per diem, and telephone and duplication charges.

I. DGS-DCLS shall derive the travel costs charged under subsections G and H of this section from the Commonwealth of Virginia reimbursement allowances and rates for lodging, per diem, and mileage. ]

1VAC30-46-160. Petitioning for a variance.

A. Any person regulated by this chapter may petition the director to grant a variance from any requirement of this chapter.  Any person submitting a petition to the director [ must shall ] meet the provisions of this section.  Any petition submitted to the director is subject to the Virginia Administrative Process Act (§2.2-4000 et seq. of the Code of Virginia).

B. The petition shall be submitted to the director by certified mail and shall include:

1. The petitioner's name and address;

2. A statement of the petitioner’s interest in the proposed action;

3. A description of desired action and a citation of the regulation from which a variance is requested;

4. A description of need and justification for the proposed action, including impact of the proposed action on the laboratory’s operation;

5. Information demonstrating that the requested variance will meet the purposes and objectives of the relevant regulatory provision and of §2.2-1105 of the Code of Virginia (Environmental Laboratory Certification Program);

6. The duration of the variance, if applicable;

7. The potential impact of the variance on public health or the environment;

8. Other information believed by the applicant to be pertinent; and

9. The following statement signed by the petitioner or authorized representative:  "I certify that I have personally examined and am familiar with the information submitted in this petition and all attached documents, and that, based on my inquiry of those individuals immediately responsible for obtaining the information, I believe that the submitted  information is true, accurate, and complete. I am aware that there are significant penalties for submitting false information, including the possibility of fine and imprisonment.”

C. Petition processing.

1. After receiving a petition that includes the information required in subsection B of this section, the director will determine whether the information received is sufficient to render the decision.  If the information is deemed insufficient, the director will specify additional information needed and request that it be furnished.

2. The petitioner may submit the additional information requested, or may attempt to show that no reasonable basis exists for additional information.  If the director agrees that no reasonable basis exists for the request for additional information, he will act in accordance with subsection D of this section.  If the director [ continues to believe finds ] that a reasonable basis exists to require the submission of such information, he will proceed with the denial action in accordance with the Administrative Process Act.

D. Public review of tentative decision.  The director will evaluate the application and issue a draft notice tentatively denying the petition, granting the variance as requested, or granting a modified or partial variance.  Notification of this tentative decision will be published in the Virginia Register of Regulations.  The director will accept comment on the tentative decision for 30 days, and shall hold a public hearing if a request is received or at his discretion if there is no request.  The director will issue a final decision after receipt of comments and after the hearing (if any).

E. Conditions for granting variance request or a modified variance.

1. The director may grant the variance if the applicant demonstrates to the satisfaction of the director that:

a. The proposed variance will meet the goals and purposes of the provisions from which a variance is sought; [ and ]

b. The variance does not conflict with federal or state law or regulations.

2. If the director grants a variance request, the notice to the petitioner shall provide that the variance may be terminated upon a finding by the director that the petitioner has failed to comply with any requirements of the variance.

3. When a modified variance is granted, the director may:

a. Specify the termination date of the variance;

b. Include a schedule for:

(1) Compliance, including increments of progress, by the laboratory with each requirement of the variance; and

(2) Implementation by the laboratory of such measures as the director finds necessary in order that the variance may be granted.

F. Decisions to grant or deny a petition [ in whole or in part, or to modify or terminate a variance ] are subject to the provisions of Article 3 (§2.2-4018 et seq.) of the Virginia Administrative Process Act.

1VAC30-46-170. (Reserved.)

1VAC30-46-180. (Reserved.)

1VAC30-46-190. (Reserved.)

Part II 
Standards

1VAC30-46-200. Incorporation of NELAC standards.

A. The [ 2002 2003 ] National Environmental Laboratory Accreditation Conference (NELAC) standards approved [ July 12, 2002 June 5, 2003 ] , as specified in 1VAC30-46-210 are incorporated by reference into this chapter.

B. Laboratories applying for accreditation and accredited under this chapter shall comply with the [ 2002 2003 ] NELAC standards incorporated by reference into 1VAC30-46-210.

[ C. The requirements of Chapter 4 of the 2003 NELAC standards, Accreditation Process, are incorporated by reference into this chapter unless the requirements are (i) already addressed in this chapter, (ii) superseded by Virginia law, or (iii) incorporated by reference into 1VAC30-46-210. ]

1VAC30-46-210. Standards for accreditation.

A. Standards for personnel qualifications. The standards for personnel qualifications are the following provisions of the National Environmental Laboratory Accreditation Conference (NELAC) standards as incorporated by reference into this part: Chapter 4, Accreditation Process, specifically, Components of Accreditation and Personnel Qualifications [ (4.1.1) and Chapter 5, Quality Systems, specifically, Technical Requirements - Personnel (5.5.2) ].

B. Standards for on-site assessment. The standards for on-site assessment are the following provisions of the NELAC standards as incorporated by reference into this part.

1. Chapter 3, On-site Assessment [ , specifically, On-site Assessment Personnel; Frequency and Types of On-site Assessments; Preassessment Procedures; Assessment Procedures; Standards for Assessment; and Documentation of On-site Assessment, ] with one exception. Subsection 3.4.5, Confidential Business Information (CBI) Considerations, shall not be incorporated by reference into this part.

2. Chapter 4, Accreditation Process, specifically, On-site Assessments and Corrective Action Reports in Response to On-site Assessment [ (4.1.2 and 4.1.3) ] .

C. Standards for proficiency testing. The standards for proficiency testing are the following provisions of the NELAC standards as incorporated by reference into this part.

1. Chapter 2, Proficiency Testing, specifically,  [ Introduction, Scope, and Applicability; ] Major PT Groups and Their Responsibilities [ - Laboratories and Accrediting Authorities ] ; Laboratory Enrollment in Proficiency Testing Programs; Requirements for Laboratory Testing of PT Study Samples; and PT Criteria for Laboratory Accreditation  [ (2.1, 2.2, 2.4 through 2.7 ].

2. Chapter 4, Accreditation Process, specifically, Proficiency Testing Samples [ (4.1.4) ] .

D. Standards for quality systems.

1. The standards for quality systems are the following provisions of the NELAC standards as incorporated by reference into this part: (i) Chapter 4, specifically, Accountability for Analytical Standards and (ii) Chapter 5, Quality Systems.

2. Quality systems - scope. Chapter 5 of the NELAC standards sets out the scope of quality systems requirements. These provisions provide an overview to major aspects of the accreditation process and are set out below for emphasis:

a. Chapter 5 includes all quality assurance policies and quality control procedures [ which that ] shall be delineated in a quality manual and followed to ensure and document the quality of the analytical data. Laboratories seeking accreditation shall assure implementation of all quality assurance policies and the essential applicable quality control procedures specified in this chapter. The quality assurance policies, which establish essential quality control procedures, are applicable to environmental laboratories regardless of size and complexity.

b. The intent of Chapter 5 is to provide sufficient detail concerning quality management requirements so that DGS-DCLS can evaluate environmental laboratories consistently and uniformly.

c. Chapter 5 sets out the general requirements that a laboratory has to successfully demonstrate to be recognized as competent to carry out specific environmental tests.

d. If more stringent standards or requirements are included in a mandated test method or by regulation, the laboratory shall demonstrate that such requirements are met. If it is not clear which [ standards or ] requirements are more stringent, the standard  [ or requirement ] from the method or regulation is to be followed.

NOTICE: The forms used in administering the above regulation are not being published; however, the name of each form is listed below. The forms are available for public inspection by contacting the agency contact for this regulation, or at the office of the Registrar of Regulations, General Assembly Building, 2nd Floor, Richmond, Virginia.

FORMS

Application for [ Accreditation Certification ] of Environmental Laboratories [ under 1VAC30-46 (with instructions), eff. xx/xx , DGS-DCLS (eff. 10/08) ].

DOCUMENTS INCORPORATED BY REFERENCE

[ 2003 ] National Environmental Laboratory Accreditation Conference [ (NELAC) ] Standards, EPA/600/R-04/003, Approved [ July 12, 2002 at Ninth NELAC Annual Meeting, June 5, 2003 ].

VA.R. Doc. No. R98-312; Filed July 16, 2008, 1:37 p.m.